## Slide 1: Biology 177: Principles of Modern Microscopy
Lecture 14:
Single Molecule Imaging
Andres Collazo, Director Biological Imaging Facility
Wan-Rong (Sandy) Wong, Graduate Student, TA

## Slide 2: Lecture 14: Single molecule imaging
Review of FRET, FLIM
Fluorescence fluctuation spectroscopy (FFS)
  Fluorescence correlation spectroscopy (FCS)
    Some concrete examples of what we can learn
  Fluorescence cross correlation spectroscopy (FCCS)
  Photon counting histogram (PCH)
FRAP/FLIP

## Slide 3
FRET:
Resonance Energy Transfer (non-radiative)
Transfer of energy from one dye to another
Depends on:
Spectral overlap
Distance
Alignment
Notes: Forster resonance energy transfer instead of fluorescence energy transfer.
Molecular yardsticks are for defining spatial dimensions

## Slide 4
RET is not always between dissimilar dyes
“Self-quenching” of dye
(“hot-potato” the energy until lost)
Depends on:
Dye Concentration
Geometry
Environment
Log [dye]
Log I
~0.1uM

## Slide 5: FRET efficiency and the Förster Equations
Distance between donor and acceptor
When r = R0, the efficiency of FRET is 50%
When R <R0, EFRET > 0.50
When R > R0, EFRET < 0.50
KT = (1/τD) • [R0/r]6
R0 = 2.11 × 10-2 • [κ2 • J(λ) • η-4 • QD]1/6
J (λ)
eA
Notes: J spectral overlap integral
R0 = the Forster radius of about 3-6 Nm
Both the rate (K(T)) and the efficiency (E(T)) of energy transfer
A plot of the FRET efficiency (EFRET) as a function of the distance (R) between a donor fluorophore (green sphere) and an acceptor fluorophore (red sphere) with an R0 of 55 Å. When R <R0, EFRET > 0.50; when R = R0, EFRET = 0.50; and when R > R0, EFRET < 0.50

## Slide 6: FRET and FLIM measure different parameters
FRET
Donor versus Acceptor fluorescent intensity
FLIM
Lifetime of donor with or without acceptor present.
Isolated donor
Donor distance too great
Donor distance correct
Notes: where τ’(DA) is the donor lifetime in the presence of the acceptor and τ(D) is the donor lifetime in the absence of the acceptor. Therefore, by measuring the donor fluorescence lifetime in the presence and absence of an acceptor (which is indicative of the extent of donor quenching due to the acceptor), it is possible to determine the distance separating donor and acceptor molecules.
Applications of fluorescence lifetime imaging (FLIM) exploit the fact that the fluorescence lifetime of a fluorophore depends on its molecular environment but not on the concentration. By using the fluorescence lifetime, molecular effects can thus be investigated independently of the unknown and usually variable fluorophore concentration.
Becker, W., Su, B., Holub, O., weisshart, K., 2011. FLIM and FCS detection in laser-scanning microscopes: Increased efficiency by GaAsP hybrid detectors. Microscopy Research and Technique 74, 804-811.

## Slide 7: Fluorescence correlation spectroscopy (FCS)
In 1972 Watt Webb’s laboratory at Cornell put fluorescence microscopy to new use
Studied reaction kinetics
Ethidium bromide binding to DNA
  Individually don’t fluoresce but together glow under UV
Could detect single molecules but could not repeatedly detect the same molecule
Notes: http://www.nature.com/milestones/milelight/full/milelight10.html
1990 invented two photon confocal microscopy

## Slide 8: Fluorescence Fluctuation Spectroscopy (FFS)
Fluorescence Correlation Spectroscopy (FCS)
Photon Counting Histogram (PCH)
Fluorescence Cross-Correlation Spectroscopy (FCCS)
  FCS with more than 1 color

## Slide 9: Fluorescence Fluctuation Spectroscopy (FFS)
Causes of fluctuations
Diffusion of labeled molecules due to Brownian motion
  In cells wide range of things cause movement (cellular trafficking, protein interaction etc.)
Photophysical processes of labeled molecules

## Slide 10: Fluctuations Carry the Information
Measured intensity fluctuations reflects (mobile fraction only)
  Number of particles
    concentration
  Diffusion of particles
    interaction
  Brightness
    Oligomerization
A particle that transits the confocal volume will generate groups of pulses.
  The correlation function calculates the mean duration time t of these groups.
  The variance/histogram of the signal yields information about oligomeric state
<I(t)>
dI(t)
dI(t)
FCS
PCH
Notes: The principle of FCS is being able to observe one or a few fluorescent particles enter and leave a small illuminated volume. If the sample, solution or cell, has a high concentration of tagged molecules it becomes hard to observe the needed fluctuations. I generally give the analogy of a crowded lecture hall. If the lecture hall is crowded and you were observing the lecture hall you would be less apt to notice if some one entered or left the room (i.e. a small fluctuation from the mean number of people in the hall). If only a few people were in the hall, then if someone left it would be very easy to notice (i.e. a large fluctuation from the mean number of people in the hall). This is why a large immobile population is not good for FCS. Fluctuations about the mean intensity is the information used for all FFS techniques.

## Slide 11: Fluorescence Fluctuation Spectroscopy (FFS)
Bacia et al., Nature Methods 2006
Notes: Bacia, K., Kim, S.A., Schwille, P., 2006. Fluorescence cross-correlation spectroscopy in living cells. Nat Methods 3, 83-89.

## Slide 12: Creating the Autocorrelation Function
“Copy” signal
Photon Burst
dI(t)
dI(t+t)
t=0
t=tD
t=inf
Notes: The fluctuations generated can be analyzed in many ways. FCS relies on the auto-correlation function which is a signal processing technique. The autocorrelation compares a signal against itself. In this case, a particle enters the focal volume and generates a photon burst. The collected signal is “copied” and the two copies are compared against one another as you shift the “copy” in time. The copy signal is shifted by some amount of time given by the greek symbol tau. All data analysis will be done in reference to the shift time tau.

## Slide 13: FCS Correlation Function
The correlation function
CF G(t)
amplitude: number of molecules
Decay time: diffusion time
offset: very slow processes

## Slide 14: Autocorrelation Function
kQ = quantum yield and detector sensitivity (how bright is our probe). This term could contain the fluctuation of the fluorescence intensity due to internal processes
W(r) describes our
observation volume
C(r,t) is a function of the fluorophore concentration over time. This is the term that contains the “physics” of the diffusion processes
Factors influencing the fluorescence signal:
Notes: Lecture 5 FCS, Autocorrelation, PCH, Cross-correlation Enrico Gratton

## Slide 15: Autocorrelation Yields Diffusion and Concentration
Fit Autocorrelation curve for Diffusion time (tD) and particle concentration N
Notes: The larger the particle, the longer it will take to transverse the focal volume thus the photon burst will be longer in time. This results in the correlation function moving the left for larger particles. The diffusion time gives information about the diffusion coefficient, the molecular weight and possibly the viscosity of the media. The correlation curve amplitude is inversely proportional to the particle concentration. The concentration, N, is the average number of molecules in the focal volume at one time. Standard dye solutions with known diffusion coefficients can be used to characterize the focal volume. Once that is done, you can use the focal volume information and the Number of particles to get an accurate measure of the MOBILE concentration (mol/L). The curve amplitude being inversely proportional to the particle number makes sense because the higher the N, the smaller the fluctuations will be from the mean intensity giving less correlation.
ALL DATA IS FIT USING A NON-LINEAR LEAST SQUARES ERRORS FUNCTION TO GET DIFFUSION TIMES AND N. You will run the risk of getting bad results with bad curve fitting or curve fitting interpretation.

## Slide 16: Autocorrelation Yields Diffusion and Concentration
Fit Autocorrelation curve for Diffusion time (tD) and particle concentration N
Notes: The larger the particle, the longer it will take to transverse the focal volume thus the photon burst will be longer in time. This results in the correlation function moving the left for larger particles. The diffusion time gives information about the diffusion coefficient, the molecular weight and possibly the viscosity of the media. The correlation curve amplitude is inversely proportional to the particle concentration. The concentration, N, is the average number of molecules in the focal volume at one time. Standard dye solutions with known diffusion coefficients can be used to characterize the focal volume. Once that is done, you can use the focal volume information and the Number of particles to get an accurate measure of the MOBILE concentration (mol/L). The curve amplitude being inversely proportional to the particle number makes sense because the higher the N, the smaller the fluctuations will be from the mean intensity giving less correlation.
ALL DATA IS FIT USING A NON-LINEAR LEAST SQUARES ERRORS FUNCTION TO GET DIFFUSION TIMES AND N. You will run the risk of getting bad results with bad curve fitting or curve fitting interpretation.

## Slide 17: Autocorrelation Yields Diffusion and Concentration
Fit Autocorrelation curve for Diffusion time (tD) and particle concentration N
Notes: The larger the particle, the longer it will take to transverse the focal volume thus the photon burst will be longer in time. This results in the correlation function moving the left for larger particles. The diffusion time gives information about the diffusion coefficient, the molecular weight and possibly the viscosity of the media. The correlation curve amplitude is inversely proportional to the particle concentration. The concentration, N, is the average number of molecules in the focal volume at one time. Standard dye solutions with known diffusion coefficients can be used to characterize the focal volume. Once that is done, you can use the focal volume information and the Number of particles to get an accurate measure of the MOBILE concentration (mol/L). The curve amplitude being inversely proportional to the particle number makes sense because the higher the N, the smaller the fluctuations will be from the mean intensity giving less correlation.
ALL DATA IS FIT USING A NON-LINEAR LEAST SQUARES ERRORS FUNCTION TO GET DIFFUSION TIMES AND N. You will run the risk of getting bad results with bad curve fitting or curve fitting interpretation.

## Slide 18: What about the excitation (or observation) volume shape?

## Slide 19: Effect of Shape on the (Two-Photon) Autocorrelation Functions:
For a 2-dimensional Gaussian excitation volume:
For a 3-dimensional Gaussian excitation volume:
1-photon equation contains a 4, instead of 8

## Slide 20: Independent Processes Contribute Fluctuations
Contributions of single independent processes multiply
More process system
1E-6
1E-5
1E-4
1E-3
0.01
0.1
1
10
100
1000
10000
1.0
1.1
1.2
1.3
1.4
1.5
G(
t
)
t
[ms]
diffusion
triplet
exponential

## Slide 21: Additional Equations for these independent processes:
... where N is the average particle number, tD is the diffusion time (related to D, tD=w2/8D, for two photon and tD=w2/4D for 1-photon excitation), and S is a shape parameter, equivalent to w/z in the previous equations.
3D Gaussian Confocor analysis:
Triplet state term:
..where T is the triplet state amplitude and tT is the triplet lifetime.

## Slide 22: Fitting with Correct Model

## Slide 23: Fitting with Correct Model
Schwille and Haustein 2004

## Slide 24: Work Flow for FCS
<I>
I(t)
t
dI(t)
1
2
3
Principle steps
Measuring fluctuation intensities
Calculating correlation function
Fitting to biophysical model
AC: compare signal w/ itself
CC: compare signal w/ another
wr2
4td,i
D=
Diffusion coefficient:

## Slide 25: FCS also benefits from FLIM
FCS measurements at single point allow kinetic and diffusion properties, concentration and aggregation state of fluorescently labeled molecules to be determined.
FLIM measurement of fluorescent lifetime of fluorophore is sensitive to the molecular environment of that fluorophore.
FCS and FLIM allow information to be gathered on diffusional mobility, protein clustering and interactions, and molecular environment.
Notes: Breusegem, S.Y., Levi, M., Barry, N.P., 2006. Fluorescence Correlation Spectroscopy and Fluorescence Lifetime Imaging Microscopy. Nephron Experimental Nephrology 103, e41-e49.

## Slide 26: Zeiss ConfoCor3: FCS Setup on a Laser Scanning Confocal Microscope
Schwille and Haustein 2004
Avalanche Photodiode Detector (APD)
Single Photon Sensitivity
Focus to tiny volume (<1 femtoliter)
Notes: http://hamamatsu.magnet.fsu.edu/articles/avalanche.html
http://www.photonics.com/EDU/Handbook.aspx?AID=25535

## Slide 27: Fluorescence Fluctuation Spectroscopy (FFS)
Bacia et al., Nature Methods 2006
wr2
4td,i
D=
Diffusion coefficient:
Notes: Bacia, K., Kim, S.A., Schwille, P., 2006. Fluorescence cross-correlation spectroscopy in living cells. Nat Methods 3, 83-89.

## Slide 28: Fluorescent recovery after photobleaching (FRAP)
Like FCS also used for calculating diffusion
Fluorescence Loss in Photobleaching (FLIP) related technique
Primary use of FLIP is to determine continuity of membranous organelles.
wr2
4td,i
D=
Diffusion coefficient:
Notes: The major difference between these two microscopy techniques is that FRAP involves the study of a cell’s ability to recover after a single photobleaching event whereas FLIP involves the study of how the loss of fluorescence spreads throughout the cell after multiple photobleaching events. This difference in purpose also leads to a difference in what parts of the cell are observed. In FRAP, the area that is actually photobleached is the area of interest. Conversely, in FLIP, the region of interest is just outside the region that is being photobleached. Another important difference is that in FRAP, there is a single photobleaching event and a recovery period to observe how well fluorophores move back to the bleached site. However, in FLIP, multiple photobleaching events occur to prevent the return of unbleached fluorophores to the bleaching region.
http://en.wikipedia.org/wiki/Fluorescence_loss_in_photobleaching
http://en.wikipedia.org/wiki/Fluorescence_recovery_after_photobleaching
http://www.olympusconfocal.com/applications/flipandfrap.html
"Frap diagram" by MDougM - Own work. Licensed under Public Domain via Wikimedia Commons - http://commons.wikimedia.org/wiki/File:Frap_diagram.svg#mediaviewer/File:Frap_diagram.svg
"Fluorescence Loss in Photobleaching Schematic" by Steven Lentz - was sent to me personally. Licensed under CC BY-SA 3.0 via Wikimedia Commons - http://commons.wikimedia.org/wiki/File:Fluorescence_Loss_in_Photobleaching_Schematic.jpg#mediaviewer/File:Fluorescence_Loss_in_Photobleaching_Schematic.jpg

## Slide 29: Two-Photon microscopy: Good for FRAP?
Optical sectioning by non-linear absorbance
Notes: While FRAP has been around for a while what development really made this technique take off?

## Slide 30: FRAP has been around since 1976
Again developed by Watt Webb’s laboratory at Cornell
Though it’s been around for a while can you guess what development really made FRAP take off?
Notes: http://www.ncbi.nlm.nih.gov/pmc/articles/PMC1334945/

## Slide 31: Green Fluorescent Protein
From marine invertebrate, Crystal jelly (Aequorea Victoria)
Can be coded in genes and made by the organism

## Slide 32: Choosing Fluorescent Proteins
Shaner, N.C., Steinbach, P.A., Tsien, R.Y., 2005. A guide to choosing fluorescent proteins. Nat Meth 2, 905-909.
Notes: Shaner, N.C., Steinbach, P.A., Tsien, R.Y., 2005. A guide to choosing fluorescent proteins. Nat Meth 2, 905-909.

## Slide 33: Fluorescence Fluctuation Spectroscopy (FFS)
Fluorescence Correlation Spectroscopy (FCS)
Photon Counting Histogram (PCH)
Fluorescence Cross-Correlation Spectroscopy (FCCS)
  FCS with more than 1 color
Advantages over FRAP and FRET

## Slide 34: Fluorescence Fluctuation Spectroscopy (FFS)
Advantages over FRAP and FRET
Table: http://www.fcsxpert.com/

## Slide 35: But what do you do when FRAP and FCS give different results?
Drosophila bicoid protein gradient
With FRAP measured D=0.3 μm2/second
With FCS measured D=∼7 μm2/second
Which result do you believe?
Notes: http://www.sciencedirect.com/science/article/pii/S0092867407006630
Wieschaus Cell paper
http://www.ncbi.nlm.nih.gov/pmc/articles/PMC2920644/
FCS paper
Development paper by Wieschaus reviewing his result
http://dev.biologists.org/content/137/14/2253.full
Schematic of the Bicoid gradient in the syncytial Drosophila embryo. (A-G) Bicoid, as a transcription factor, accumulates within nuclei during interphase. The nuclear Bicoid gradient, which decreases in concentration with distance from the anterior pole, forms in less than 3 hours in a Drosophila embryo that is ~500 μm long (AP axis) and can be observed (B) before nuclei migrate to the surface of the embryo. During the first ~80 minutes (corresponding to cell cycles 1-8), nuclei proliferate in the centre of the embryo (A-C), and then migrate to the surface of the embryo in cycle 9. During the last five cell cycles, nuclei divide at the surface of the embryo (D,E). Nuclei replicate every ~9 minutes during the first nine cell cycles, but this slows to ~20 minutes in cycle 13. Nuclei are ~10 μm in diameter in pre-cycle 14 embryos (A-D) but ~6 μm in cycle 14 (E). The extent to which the mRNA encoding Bicoid, which is initially tightly localised to the anterior pole (A), becomes delocalised is a matter of controversy. Two alternative bicoid RNA distributions during cycle 14 are shown in E and F. (G) Illustration of the nucleus and surrounding cytoplasm. (H) In species that produce larger or smaller eggs than Drosophila melanogaster, the Bicoid gradient scales with egg length.

## Slide 36: Meeting Summary on Impact of Genomics on Drug Discovery and Development
Rastan, S. (2001). Genomics: saviour or millstone? Trends in Genetics 17, 247-248.
Genome (What could happen)
Transcriptome (What might be happening)
Proteome (What is happening)
Lucitt, M. B., et al. (2008). Analysis of the Zebrafish Proteome during Embryonic Development. Molecular & Cellular Proteomics 7, 981-994.
Levi, L., et al. (2009). Revealing genes associated with vitellogenesis in the liver of the zebrafish (Danio rerio) by transcriptome profiling. BMC Genomics 10, 141.

## Slide 37: Flip trap screen (http://www.fliptrap.org) Le Trinh et al. Gene Dev. 2011
Gene trapping vector: Citrine (YFP) flanked by splice acceptor & donor, forward orientation; mCherry (RFP) polyadenylation signal, reverse orientation; lox & FRT sites
Transposon based gene trapping technology (Tol2)
Notes: Visualize proteome via fluorescent tagging of full-length endogenous proteins
Flip trap screen done in the laboratories of Scott Fraser and Marianne Bronner at Caltech published in 2011

## Slide 38: Flip Trap Screen Labels Endogenous Proteins: Different Sub-Cellular Compartments and Cell Types
Zebrafish Ear Development

## Slide 39: FCS Detects Intracellular Variations

## Slide 40: Diffusion Coefficients Similar to those Obtained Using FRAP

## Slide 41: Epithelial transitions during pharyngeal pouch formation
Wnt signaling role in pharyngeal pouch formation
Adherens junction
Choe Chong P, Collazo A, Trinh Le A, Pan L, Moens Cecilia B, Crump JG (2013) Wnt-Dependent Epithelial Transitions Drive Pharyngeal Pouch Formation. Developmental Cell 24: 296-309.

## Slide 42: Alcama immunoglobulin-domain protein functions to restabilize adherens junctions
Use α-Catenin protein labeled transgenic zebrafish line to study protein mobility at adherens junction.
Alcama associates with α-Catenin at adherens junction
At transitional stage endodermal epithelium is destabilized

## Slide 43: Published FCS Studies in Zebrafish Have Used Exogenously Expressed Proteins
Proteins acting as morphogen
  Yu, S. R., M. Burkhardt, et al. (2009). Nature 461(7263): 533-536.
Receptor ligand interactions
  Ries, J., S. R. Yu, et al. (2009). Nat Methods 6(9): 643-645.

## Slide 44: Less Gnb2 in Hair Cells than Neurons

## Slide 45: Expressing Tagged Proteins by mRNA Injection Can Lead to Over-expression and Incorrect Estimates of Diffusion Rates

## Slide 46: What Do Differences in Diffusion Coefficients Mean?
Binding to larger protein?

## Slide 47: What Do Differences in Diffusion Coefficients Mean?
Binding to larger protein?
  But 100x the mass only 1/3 reduction.

## Slide 48: What Do Differences in Diffusion Coefficients Mean?
Binding to larger protein?
  But 100x the mass only 1/3 reduction.
More likely causing interactions with cytoskeleton, organelles or other structures
Nucleus
Cell Membrane
Mitochondria

## Slide 49: Conclusion of examples
FCS + Flip Trap lines provide unique insights into protein kinetics in vivo.
    Intracellular differences across subcellular regions
    Developmental transitions in protein kinetics
Endogenous protein concentrations crucial for accurate FCS measurements

## Slide 50: Fluorescence Cross-Correlation Spectroscopy
Cross-Correlation uses spectrally separable fluorophores to probe for interaction
Cross Correlation Curve Amplitude directly relates to interaction
Photon Burst Ch.1
Photon Burst Ch.2
Notes: FCCS is not too much different than single color FCS. FCCS is used to observe interaction and is much more sensitive that using differing diffusion times in single color FCS. Instead of making a “copy” signal you have two signals. The signals are then correlated together and the resulting curve gives you information about particle interaction.

## Slide 51: Cross Correlation Reveals Details of Particle Binding
Autocorrelation reveals portion of unbound particles
Cross correlation reveals bound portion
Binding constant can be calculated

## Slide 52: Controls Provide a Basis for Comparison for Cross-Correlation
Spectral cross talk can lead to false cross-correlation
With dual excitation lasers and differing expression levels, cross-correlation is never “perfect”
Controls needed for Comparison
Slaughter et al, PNAS. 2007.
Notes: Some practical considerations. Need good spectral separation, eGFP-mCherry work fine, but have some bleed through. Too much bleed through can give a high amount of false cross-correlation. With 2-color one photon FCCS the overlap for both excitation lasers is not perfect. That is to say, a 488 laser will have a smaller PSF than the 561nm laser. This results in cross-correlation standards giving results that are not “perfect.” If perfect the cross correlation curve will overlap the two auto-correlation curves. Need controls!
http://www.pnas.org/content/104/51/20320.full

## Slide 53: This is same issue we saw with FRET
Notes: Figure 1. Caspase-3 biosensors based on dual FRET pairs. (a,b) Schematics of caspase-3 biosensors. "DEVD" represents the sequence LGGTGSGSGDEVDG. Numbers indicate first and last residue of each fluorescent protein. (c) The emission spectrum of mAmetrine-DEVD-tdTomato before and after proteolysis, the excitation spectrum of mAmetrine, and the transmission profiles of excitation and emission filters used for FRET imaging. (d) The emission spectrum of mCitrine-DEVD-mTFP1 before and after proteolysis, the excitation spectrum of mTFP1, and the profiles of the excitation and emission filters.
Ai, H.-w., Hazelwood, K.L., Davidson, M.W., Campbell, R.E., 2008. Fluorescent protein FRET pairs for ratiometric imaging of dual biosensors. Nat Meth 5, 401-403.

## Slide 54: Photon Counting Histogram
Photon count distribution originates from a convolution of two sources
  Photon Detection Statistics
    Poisson
  Particle Number Fluctuations
    Poisson
Further complications come from variations in PSF
Information Gained
  Concentration
  Brightness
(Qian and Elson, 1989, Applied Polymer Symposia. John Wiley and Sons, New York. 305-314,1990; Chen et al, 1999 Biophys J. 77: 553–567.; Muller et al, 2000 Biophys J 78:474–486).
Photon Counts
frequency
(en=1.0)
(en=2.2)
(en=3.7)
Increasing Brightness

## Slide 55: Brightness Reveals Oligomerization
Slaughter et al, PloS ONE. 2008.

## Slide 56: Homework 5
Early on during zebrafish development, many molecules are involved in patterning the embryo’s tissues and axes. One possible explanation for this complex patterning is Alan Turing’s reaction-diffusion model. To test this you have fluorescently tagged three proteins involved in this process. Your hypothesis is that one protein acts at a long distance while another acts at a short distance and the third at an intermediate distance.
Questions: What fluorescent technique would you use to determine the mobility of these three proteins? What would you predict is the relative mobility of these three proteins?

## Slide 57: Paper to read
Pearson, H., 2007. The good, the bad and the ugly. Nature 447, 138-140.
http://www.nature.com/nature/journal/v447/n7141/full/447138a.html

## Slide 58: Cautionary tale of fluorescence microscopy

## Slide 59: Cautionary tale of fluorescence microscopy

## Slide 60: Choice quotes
The modern light microscope comes with the accoutrements and price tag of a high-speed racing car and offers an exhilarating ride…But not everyone should be allowed behind the eyepiece.
She aims to scare users enough that they will consult her before embarking on a doomed microscopy project: “It’s quite cruel of me isn’t it?”
It’s so temperamental that Waters says she advises new graduate students “to turn around and run away” if a prospective supervisor suggests FRET for their thesis work.

## Slide 61: Top tips for taking images
Choose the right method
Prepare sample carefully
Choose the right mountant
Select objective lens with care
Choose right tags and filters
Avoid aberration
Don’t saturate the image
Don’t always select the ideal cell
Keep your cells happy
Notes: Good advice

## Slide 62: Multispectral FRET is more sensitive
Multispectral techniques can extract even weak FRET signal
Changes of this size are the break-even point for barrier filters vs. single spectral channel