## Slide 1: Bi/BE 177: Principles of Modern Microscopy
Lecture 15: FRET, TIRFM, Super-resolution microscopy Part I
Andres Collazo, Director Biological Imaging Facility
Ke Ding, Graduate Student, TA
Wan-Rong (Sandy) Wong, Graduate Student, TA

## Slide 2: Lecture 15: FRET, TIRF, NSOM
FRET
FLIM review
Total internal reflection fluorescence microscopy (TIRFM)
Super resolution microscopy
  NSOM
Critiquing figures

## Slide 3: Questions about last lecture?

## Slide 4: Förster Resonance Energy Transfer (FRET)
Great method for the detection of:
  Protein-protein interactions
  Enzymatic activity
  Small molecules interacting inside a cell
Notes: http://www.nature.com/nmeth/journal/v5/n5/full/nmeth.1207.html

## Slide 5
If dye at high concentration
“hot-potato” the energy until lost
Remember our fluorescence discussion?
Resonance Energy Transfer (non-radiative)
The Bad: Self-quenching

## Slide 6
Resonance Energy Transfer (non-radiative)
“Self-quenching” of dye
(“hot-potato” the energy until lost)
Depends on:
Dye Concentration
Geometry
Environment
Log [dye]
Log I
~0.1uM

## Slide 7
FRET:
Resonance Energy Transfer (non-radiative)
The Good: FRET as a molecular yardstick
Transfer of energy from one dye to another
Depends on:
Spectral overlap
Distance
Alignment
Notes: Forster resonance energy transfer instead of fluorescence energy transfer.
Molecular yardsticks are for defining spatial dimensions

## Slide 8
FRET:
Optimize spectral overlap
Optimize k2 -- alignment of dipoles
Minimize direct excitement of the acceptor
(extra challenge for filter design)
donor
acceptor
Notes: Perpendicular is not good.

## Slide 9: FRET Diagram
4nsec
0.8 emitted
Non-radiative transfer
-xx-
Less
-xx-
Less

## Slide 10
The Förster Equations.
r is the center-to-center distance (in cm) between the donor and acceptor
tD is the fluorescence lifetime of the donor in the absence of FRET
k2 is the dipole-dipole orientation factor,
QD is the quantum yield of the donor in the absence of the acceptor
 is the refractive index of the intervening medium,
FD (l) is the fluorescence emission intensity at a given wavelength l (in cm)
eA (l) is the extinction coefficient of the acceptor (in cm -1 M -1).
The orientation factor k2 can vary between 0 and 4, but
typically k2 = 2/3 for randomly oriented molecules (Stryer, 1978).
When r = R0, the efficiency of FRET is 50%
(fluorescein-tetramethylrhodamine pair is 55 Å)
KT = (1/τD) • [R0/r]6
R0 = 2.11 × 10-2 • [κ2 • J(λ) • η-4 • QD]1/6
J (λ)
eA
Notes: J spectral overlap integral
R0 = the Forster radius of about 3-6 Nm
E = the FRET energy transfer efficiency
Both the rate (K(T)) and the efficiency (E(T)) of energy transfer are directly related to the lifetime of the donor fluorophore in the presence and absence of the acceptor.
http://www.olympusmicro.com/primer/techniques/fluorescence/fret/fretintro.html
http://en.wikipedia.org/wiki/F%C3%B6rster_resonance_energy_transfer
http://www.nature.com/nbt/journal/v21/n11/full/nbt896.html

## Slide 11
FRET:
Resonance Energy Transfer (non-radiative)
The Good: FRET as a molecular yardstick
Transfer of energy from one dye to another
Depends on:
Spectral overlap
Distance
Alignment
Notes: Forster resonance energy transfer instead of fluorescence energy transfer.
Molecular yardsticks are for defining spatial dimensions

## Slide 12
Remember:
Going back to our Fluorescence lecture
How dipole affects FRET as a molecular yardstick
Fluorescent Dye
Dipole antenna
Delocalized electrons
Longer dipole, longer λ
Notes: transition dipole moment (TDM)
http://www.nh.cas.cz/people/lazar/celler/emission.php

## Slide 13: Fluorescent dye as dipole antenna
Absorption depends on orientation

E

Propagation direction
E

Propagation direction
Notes: electric dipole moment (µ) of the molecule
Quantum Yield and Polarization (1) Joachim Mueller
Fluorescence Workshop UMN Physics
June 8-10, 2006
Quantum yield, polarized light, dipole moment, photo- selection, dipole radiation, polarization and anisotropy
Antenna pictures from
http://www.nh.cas.cz/people/lazar/celler/emission.php

## Slide 14: Fluorescent dye as dipole antenna
Orientation of fluorescence emission
No emission along the dipole axis
μE
No emission along the dipole axis
Maximal emission normal to the dipole axis
Maximal emission normal to the dipole axis

Some emission along this direction
Dipole radiation pattern
Notes: electric dipole moment (µ) of the molecule
Quantum Yield and Polarization (1) Joachim Mueller
Fluorescence Workshop UMN Physics
June 8-10, 2006
Quantum yield, polarized light, dipole moment, photo- selection, dipole radiation, polarization and anisotropy
Antenna pictures from
http://www.nh.cas.cz/people/lazar/celler/emission.php

## Slide 15: Fluorescent dye as dipole antenna
Orientation of fluorescence emission affects FRET efficiency
Notes: electric dipole moment (µ) of the molecule
Quantum Yield and Polarization (1) Joachim Mueller
Fluorescence Workshop UMN Physics
June 8-10, 2006
Quantum yield, polarized light, dipole moment, photo- selection, dipole radiation, polarization and anisotropy
Antenna pictures from
http://www.nh.cas.cz/people/lazar/celler/emission.php

## Slide 16: More about FRET (Förster Resonance Energy Transfer)
Isolated donor
Effective between 10-100 Å only
Emission and excitation spectrum must significantly overlap
Note: donor transfers non-radiatively to the acceptor
Donor distance too great
Donor distance correct
From J. Paul Robinson, Purdue University

## Slide 17: FRET efficiency and the Förster Equations
Distance between donor and acceptor
When r = R0, the efficiency of FRET is 50%
When R <R0, EFRET > 0.50
When R > R0, EFRET < 0.50
KT = (1/τD) • [R0/r]6
R0 = 2.11 × 10-2 • [κ2 • J(λ) • η-4 • QD]1/6
J (λ)
eA
Notes: J spectral overlap integral
R0 = the Forster radius of about 3-6 Nm
Both the rate (K(T)) and the efficiency (E(T)) of energy transfer
A plot of the FRET efficiency (EFRET) as a function of the distance (R) between a donor fluorophore (green sphere) and an acceptor fluorophore (red sphere) with an R0 of 55 Å. When R <R0, EFRET > 0.50; when R = R0, EFRET = 0.50; and when R > R0, EFRET < 0.50

## Slide 18: Optimizing FRET: Designs of new FRET pairs
Difficult to find two FRET pairs that can use in same cell
Used as Caspase 3 biosensors and for ratiometric imaging
Notes: Figure 1. Caspase-3 biosensors based on dual FRET pairs. (a,b) Schematics of caspase-3 biosensors. "DEVD" represents the sequence LGGTGSGSGDEVDG. Numbers indicate first and last residue of each fluorescent protein. (c) The emission spectrum of mAmetrine-DEVD-tdTomato before and after proteolysis, the excitation spectrum of mAmetrine, and the transmission profiles of excitation and emission filters used for FRET imaging. (d) The emission spectrum of mCitrine-DEVD-mTFP1 before and after proteolysis, the excitation spectrum of mTFP1, and the profiles of the excitation and emission filters.
Ai, H.-w., Hazelwood, K.L., Davidson, M.W., Campbell, R.E., 2008. Fluorescent protein FRET pairs for ratiometric imaging of dual biosensors. Nat Meth 5, 401-403.

## Slide 19: Fluorophore brightness =  Q
DsRed Q ~ 0.79 x 75,000 ~ 59,250 M-1.cm-1 (100%)
mRFP1 Q ~ 0.25 x 50,000 ~ 12,500 M-1.cm-1 (21%)
eGFP Q ~ 0.6 x 55,000 ~ 33,000 M-1.cm-1 (56%)
Fluorescein Q ~ 0.8 x 70,000 ~ 56,000 M-1.cm-1 (95%) (dye!)
Notes: Shaner et al, Nature Biotechnology, 2004

## Slide 20: Optimizing FRET: Designs of new FRET pairs
mAmetrine developed by directed protein evolution from violet excitable GFP variant
Bright, extinction coefficient = 44,800 M-1 cm-1
Quantum yield = 0.58
But bleaches, 42% of mCitrine time and 1.7% of tdTomato
Notes: Ai, H.-w., Hazelwood, K.L., Davidson, M.W., Campbell, R.E., 2008. Fluorescent protein FRET pairs for ratiometric imaging of dual biosensors. Nat Meth 5, 401-403.

## Slide 21: Problems with FRET
4nsec
The acceptor excited directly by the exciting light
  “FRET” signal with no exchange
  Increased background
  Decreases effective range for FRET assay

## Slide 22: Problems with FRET
2. Hard to really serve as a molecular yardstick*
  Orientation seldom known
      assume k2 = 2/3 (random assortment)
  Exchange depends on environment of dipoles
  Amount of FRET varies with the lifetime of the donor fluorophore
* r = R0, the efficiency of FRET is 50%
(fluorescein-tetramethylrhodamine pair is 55 Å)
Notes: See this paper for more Kappa squared and other variable cautions.
http://www.nature.com/nbt/journal/v21/n11/full/nbt896.html

## Slide 23: Amount of FRET varies with the lifetime of the donor fluorophore
4nsec
Longer lifetime of the donor gives longer time to permit the energy transfer (more for longer)
Added Bonus: Allows lifetime detection to reject direct excitement of the acceptor (FRET=late)

## Slide 24: Fluorescence Lifetime Imaging Microscopy (FLIM)
Measure spatial distribution of differences in the timing of fluorescence excitation of fluorophores
Combines microscopy with fluorescence spectroscopy
Fluorescent lifetimes very short (ns) so need fast excitation and/or fast detectors
Requirements for FLIM instruments
  Excitation light intensity modulated or pulsed
  Emitted fluorescence measured time resolved
Notes: FLIM has its roots in two fields of research: 1) microscopy and 2) fluorescence spectroscopy. In the latter field of research, non-spatially-resolved fluorescence lifetime measurements were performed since 1926 [1], i.e. long before FLIM was developed. Typically, bulk measurements were carried out using cuvettes. Not surprisingly, most of the methodology and nomenclature used in FLIM today, e.g.‘frequency-domain’, and ‘time-domain’, have their origins in instruments that were used for cuvette-based lifetime measurements.
van Munster, E., Gadella, T.J., 2005. Fluorescence Lifetime Imaging Microscopy (FLIM), in: Rietdorf, J. (Ed.), Microscopy Techniques. Springer Berlin Heidelberg, pp. 143-175.

## Slide 25: Fluorescence Lifetime Imaging Microscopy (FLIM)
Two methods for FLIM
  Frequency-domain
    Intensity of excitation light continuously modulated
    For emission measure phase shift & decrease in modulation
  Time-domain
    Pulsed excitation that is faster than fluorescence lifetime
    Emission measurement is time-resolved
Notes: The first instrument combining time resolved fluorescence spectroscopy with microscopy, dates back to 1959 [2]. In this instrument, only single point measurements could be done, so strictly speaking no actual imaging was done. The first instrument measuring spatially resolved lifetimes was described in 1989 [3]. Due to the requirement of short light pulses and fast detection, time-domain measurements became possible only about 40 years later than frequency-domain measurements using a flashlamp as excitation source [16].
van Munster, E., Gadella, T.J., 2005. Fluorescence Lifetime Imaging Microscopy (FLIM), in: Rietdorf, J. (Ed.), Microscopy Techniques. Springer Berlin Heidelberg, pp. 143-175.

## Slide 26: FRET and FLIM
Donor fluorescence lifetime during FRET reduced compared to control donor fluorescence lifetime
During FRET, donor fluorescence lifetime less than control donor fluorescence lifetime (tD)
But isn’t it easier to image decreases in donor fluorescence intensity rather than measure fluorescence lifetime?
KT = (1/τD) • [R0/r]6
Notes: Both the rate (K(T)) and the efficiency (E(T)) of energy transfer are directly related to the lifetime of the donor fluorophore in the presence and absence of the acceptor.
In other words the donor fluorescence (or excited-state) lifetime in a FRET situation (tFRET) is reduced as compared to the control donor fluorescence lifetime t (see Eq. 14 in which
  tFRET = t (1 – E) (14)
where E is the energy transfer efficiency).
http://olympus.magnet.fsu.edu/primer/techniques/fluorescence/fret/fretintro.html
van Munster, E., Gadella, T.J., 2005. Fluorescence Lifetime Imaging Microscopy (FLIM), in: Rietdorf, J. (Ed.), Microscopy Techniques. Springer Berlin Heidelberg, pp. 143-175.

## Slide 27: FRET and FLIM: addressing nonlinearities
Brightness (or intensity) of fluorophore, as measured on your image, more than just Q
  Local concentration of fluorophore
  Optical path of microscope
  Local excitation light intensity
  Local fluorescence detection efficiency
FLIM provides independent measure of local donor lifetime
Notes: The problem with quantitative imaging of intensities is that they do not depend only on the local quantum yield of the fluorophores but also on i) the local concentration of the fluorophore, ii) the optical path of the micro- scope, iii) the local excitation light intensity and iv) the local fluorescence detection efficiency. Especially the local concentration of the fluorophore is not easily determined separately from the local quantum yield. In contrast, FLIM does provide an independent estimate of the local donor lifetime (proportional to the donor quantum yield), independent of i–iv mentioned above. Hereby FLIM is a robust technique for measuring FRET.
All those nonlinearities cause problems.
Q=molar absorption coefficient times quantum yield.
van Munster, E., Gadella, T.J., 2005. Fluorescence Lifetime Imaging Microscopy (FLIM), in: Rietdorf, J. (Ed.), Microscopy Techniques. Springer Berlin Heidelberg, pp. 143-175.

## Slide 28: FRET and FLIM measure different parameters
FRET
Donor versus Acceptor fluorescent intensity
FLIM
Lifetime of donor with or without acceptor present.
Isolated donor
Donor distance too great
Donor distance correct
Notes: where τ’(DA) is the donor lifetime in the presence of the acceptor and τ(D) is the donor lifetime in the absence of the acceptor. Therefore, by measuring the donor fluorescence lifetime in the presence and absence of an acceptor (which is indicative of the extent of donor quenching due to the acceptor), it is possible to determine the distance separating donor and acceptor molecules.
Applications of fluorescence lifetime imaging (FLIM) exploit the fact that the fluorescence lifetime of a fluorophore depends on its molecular environment but not on the concentration. By using the fluorescence lifetime, molecular effects can thus be investigated independently of the unknown and usually variable fluorophore concentration.
Becker, W., Su, B., Holub, O., weisshart, K., 2011. FLIM and FCS detection in laser-scanning microscopes: Increased efficiency by GaAsP hybrid detectors. Microscopy Research and Technique 74, 804-811.

## Slide 29: Going back to those problems with FRET: These drawbacks can all be used to make sensors
Change in FRET for changes in:
Orientation
  cameleon dye for Ca++
Local environment
  Phosphate near fluorophore
  Membrane voltage (flash)
Change in lifetime of donor
  Binding of molecule displacing water

## Slide 30: Cameleon: FRET-based and genetically-encoded calcium probe
Miyawaki et al, Nature, 1997
Calmodulin bonds Ca2+
and changes its conformation
[Ca2+]
Cameleon family: calmodulin-based indicators of [Ca2+] using FRET
isosbestic point
Notes: Isosbestic point best for ratiometric measures
http://www.nature.com/nature/journal/v388/n6645/full/388882a0.html

## Slide 31: Paper to read
Pearson, H., 2007. The good, the bad and the ugly. Nature 447, 138-140.
http://www.nature.com/nature/journal/v447/n7141/full/447138a.html

## Slide 32: Single molecule tracking
High speed
Single molecule imaging
  Fluorescence correlation spectroscopy (FCS)
  Total internal reflection microscopy (TIRF)
Super-resolution
qi
Interface
qi
Notes: http://www.rsc.org/ConferencesAndEvents/RSCConferences/FD/molecule-fd2015/
A practical superlens, super lens or perfect lens, is a lens which uses metamaterials to go beyond the diffraction limit. The diffraction limit is an inherent limitation in conventional optical devices or lenses.[1] In nano-optics, a plasmonic lens generally refers to a lens for surface plasmon polaritons (SPPs), i.e. a device that redirects SPPs to converge towards a single focal point. Since SPPs can have very small wavelength, they can converge into a very small and very intense spot, much smaller than the free-space wavelength and the diffraction limit.[1][2] Surface plasmon polaritons (SPPs), are infrared or visible-frequency electromagnetic waves, which travel along a metal-dielectric or metal-air interface. The term "surface plasmon polariton" explains that the wave involves both charge motion in the metal ("surface plasmon") and electromagnetic waves in the air or dielectric ("polariton").[1]
They are a type of surface wave, guided along the interface in much the same way that light can be guided by an optical fiber. SPPs are shorter in wavelength than the incident light (photons).[2]
Metamaterials are artificial materials engineered to have properties that have not yet been found in nature.

## Slide 33: Total internal reflection fluorescence (TIRF) microscopy
Technique that dominates most single molecule imaging approaches
Notes: http://micro.magnet.fsu.edu/primer/techniques/fluorescence/tirf/tirfhome.html
http://www.microscopyu.com/articles/fluorescence/tirf/tirfintro.html
http://olympus.magnet.fsu.edu/primer/techniques/fluorescence/tirf/tirfhome.html
http://www.nature.com/milestones/milelight/full/milelight17.html

## Slide 34
sin q critical = h1 / h2
Internal reflection depends on refractive index differences
Notes: Total internal reflection does not occur suddenly as a new phenomenon at the critical angle, but a continuous transition is followed from predominant refraction with a small amount of reflection, to total reflection when the critical angle is exceeded.
Matching refractive indices can eliminate internal reflection. Internal reflection depends on refractive index differences

## Slide 35: Evanescent waves
Near-field phenomenon
Higher frequency, more information
Formed at boundary between two media with different wave motion properties
Evanescent waves quantum tunneling phenomenon
Product of Schrödinger wave equations
Exponential decay
Notes: http://www.acs.psu.edu/drussell/Demos/EvanescentWaves/EvanescentWaves.html

## Slide 36: Metamaterials with negative refractive indices could be used to make superlenses for super resolution microcopy
Maxwell's fish-eye lens could do it with positive refractive indices
  Refractive index changes across lens (blue shading)
Harness information on resolution from evanescent waves
Type of Luneburg lens
Tyc T, Zhang X (2011) Forum Optics: Perfect lenses in focus. Nature 480: 42-43.
Notes: Darker blue, higher refractive index. Like a gravitational lens
A practical superlens, super lens or perfect lens, is a lens which uses metamaterials to go beyond the diffraction limit. The diffraction limit is an inherent limitation in conventional optical devices or lenses.[1] In nano-optics, a plasmonic lens generally refers to a lens for surface plasmon polaritons (SPPs), i.e. a device that redirects SPPs to converge towards a single focal point. Since SPPs can have very small wavelength, they can converge into a very small and very intense spot, much smaller than the free-space wavelength and the diffraction limit.[1][2] Surface plasmon polaritons (SPPs), are infrared or visible-frequency electromagnetic waves, which travel along a metal-dielectric or metal-air interface. The term "surface plasmon polariton" explains that the wave involves both charge motion in the metal ("surface plasmon") and electromagnetic waves in the air or dielectric ("polariton").[1]
They are a type of surface wave, guided along the interface in much the same way that light can be guided by an optical fiber. SPPs are shorter in wavelength than the incident light (photons).[2]
Metamaterials are artificial materials engineered to have properties that have not yet been found in nature.
https://en.wikipedia.org/wiki/Luneburg_lens

## Slide 37: TIRFM illumination configurations
Prism method
Objective Lens method
Ideally NA of 1.45 or higher
Notes: Combining this relationship with the condition for total internal reflection given above illustrates that living cells having a typical refractive index of 1.38 require illumination with an objective having a numerical aperture of greater than 1.38 in order to achieve total internal reflection. Light entering the objective must pass through the portion of the aperture cone corresponding to numerical aperture values larger than 1.38 in order to be totally reflected at the specimen-glass interface. If coherent laser illumination is employed, it must be focused at the periphery of the objective rear aperture to ensure that light will exit the front optical surface at an angle equal to or greater than the critical value. In the case of non-coherent illumination, such as that from an arc-discharge lamp, a mask in the form of an opaque disk must be introduced into the optical path to restrict light passing through the objective to the outer region of the rear aperture.
http://www.microscopyu.com/articles/fluorescence/tirf/tirfintro.html

## Slide 38: TIRFM illumination configurations
Prism method
Restricts access to specimen (difficult to manipulate)
Most illuminate opposite objective so have to pass through specimen
If prism on same side then more complicated alignment
Objective Lens method
This is the way to go
But …
Notes: But… less clean evanescent wave due to increased stray light.

## Slide 39: TIRFM applications
Benefits for imaging minute structures or single molecules in specimens with tons of fluorescence outside of optical plane of interest
Examples: Brownian motion of molecules in solution, vesicles undergoing endocytosis or exocytosis, or single protein trafficking in cells
Can get dramatic increase in signal-to-noise ratio from thin excitation region
Microsphere example
Notes: Figure 4 presents images acquired of a solution of fluorescent microspheres utilizing the TIRFM method (Figure 4(b)) and conventional epi-fluorescence illumination (Figure 4(d)). To the left of each image is its corresponding intensity histogram (Figures 4(a) and 4(c)). The improved resolution of the spheres afforded by increasing the signal-to-noise ratio (S/N) from 1.3 to 35 is apparent in the images, and in the sharp localization and higher signal intensity in the histogram corresponding to the TIRFM image (Figure 4(a)).
http://www.microscopyu.com/articles/fluorescence/tirf/tirfintro.html

## Slide 40: TIRFM applications
Ideal tool for investigation of both the mechanisms and dynamics of many of the proteins involved in cell-cell interactions
Live cell imaging
GFP-vinculin to see focal adhesions on coverslip
Notes: Figure 5 presents comparative images of live cells (PtK1 kangaroo kidney epithelial cells expressing GFP-vinculin) utilizing a conventional widefield epi-fluorescence method (Figure 5(a)) and evanescent wave illumination (Figure 5(b)). The TIRFM image reveals localization of the fusion protein in cell focal adhesions at the substrate interface in dramatic contrast to the blur produced by out-of-plane fluorescence in the epi-illumination image. Live-cell imaging represents one of the most promising applications of the TIRFM technique.

## Slide 41: TIRFM applications
Single molecule imaging
Time lapse of GFP-Rac moving along filopodia
In fact, most single molecule imaging today done with TIRFM
Notes: At the biomolecular level, TIRFM techniques have been utilized to image single molecules of the mutant protein GFP-Rac trafficking along thin filopodia of cells growing on a substrate (Figure 6). This protein is involved in cell motility, and knowledge of the dynamics of its interactions at the cell membrane are crucial to understanding the process. The visualization of single-molecule fluorescence with sufficient temporal resolution for dynamic studies is possible with TIRFM because of the outstanding signal-to-noise ratio afforded by the evanescent wave excitation. Figure 6 presents four sequential time lapse frames taken at 200-millisecond intervals, illustrating the movement of a GFP-Rac fusion protein molecule (arrows) through a fine filopodium of a Xenopus cell growing out on a substrate.
http://www.microscopyu.com/articles/fluorescence/tirf/tirfintro.html

## Slide 42: TIRFM versus Confocal Microscopy
Confocal not limited to plane at interface, can go deeper
TIRFM has thinner optical section (100 nm vs 600 nm)
TIRFM, like two photon, only excites sample at focal plane
TIRFM is cheaper to implement than confocal
TIRFM is NOT super-resolution (except in Z)
Notes: http://olympus.magnet.fsu.edu/primer/techniques/fluorescence/tirf/tirfintro.html

## Slide 43: Highly inclined and laminated optical sheet (HILO) microscopy
How to make TIRF microscope go deeper
Use a highly inclined thin illumination
Like TIRF a wide field technique
Notes: Tokunaga, M., Imamoto, N., Sakata-Sogawa, K., 2008. Highly inclined thin illumination enables clear single-molecule imaging in cells. Nat Methods 5, 159-161.

## Slide 44: Spatial Resolution of Biological Imaging Techniques
Notes: http://zeiss-campus.magnet.fsu.edu/articles/superresolution/introduction.html

## Slide 45: Super-resolution microscopy
“True” super-resolution techniques
  Subwavelength imaging
  Capture information in evanescent waves
  Quantum mechanical phenomenon
“Functional” super-resolution techniques
  Deterministic
    Exploit nonlinear responses of fluorophores
  Stochastic
    Exploit the complex temporal behaviors of fluorophores
Notes: They fall into two broad categories, "true" super-resolution techniques, which capture information contained in evanescent waves, and "functional" super-resolution techniques, which use clever experimental techniques and known limitations on the matter being imaged to reconstruct a super-resolution image.[2]
True subwavelength imaging techniques include those that utilize the Pendry Superlens and near field scanning optical microscopy, the 4Pi Microscope and structured illumination microscopy technologies like SIM and SMI. However, the majority of techniques of importance in biological imaging fall into the functional category.
There are two major groups of methods for functional super-resolution microscopy:
Deterministic super-resolution: The most commonly used emitters in biological microscopy, fluorophores, show a nonlinear response to excitation, and this nonlinear response can be exploited to enhance resolution. These methods include STED, GSD, RESOLFT and SSIM.
Stochastic super-resolution: The chemical complexity of many molecular light sources gives them a complex temporal behaviour, which can be used to make several close-by fluorophores emit light at separate times and thereby become resolvable in time. These methods include SOFI and all single-molecule localization methods (SMLM) such as SPDM, SPDMphymod, PALM, FPALM, STORM and dSTORM.

## Slide 46: Spatial Resolution of Biological Imaging Techniques
“True” super-resolution
“Functional”
Notes: http://zeiss-campus.magnet.fsu.edu/articles/superresolution/introduction.html

## Slide 47: Remember the different types of microscopy from previous lecture?
Wide-field microscopy
  Illuminating whole field of view
Confocal microscopy
  Spot scanning
Near-field microscopy
  For super-resolution

## Slide 48: Near-Field Scanning Optical Microscopy (NSOM)
Scanning Near-Field Optical Microscopy (SNOM)
Likely the super-resolution technique with the highest resolution
But only for superficial structures
A form of Scanning Probe Microscopy (SPM)

## Slide 49: All the types of microscopes
Notes: https://en.wikipedia.org/wiki/Microscope
http://micro.magnet.fsu.edu/primer/techniques/nearfield/nearfieldintro.html
Figure 7

## Slide 50: Near-Field Scanning Optical Microscopy (NSOM) Break the diffraction limit by working in the near-field
Launch light through small aperture
Illuminated “spot” is smaller than diffraction limit
(about the size of the tip for a distance equivalent to tip diameter)
Near-field = distance of a couple of tip diameters
Notes: One of the more intuitive super-resolution techniques

## Slide 51: NSOM working in the near-field
Aperture diameter less than the wavelength of light
In 1993 Eric Betzig and Robert Chichester used NSOM for repetitive single molecule imaging
Notes: In 1993, this situation changed markedly when Eric Betzig and Robert Chichester reported the first repetitive imaging of single fluorophores at room temperature with a new technique called near-field scanning optical microscopy (NSOM) that repeatedly scans an extremely small optical probe over a sample. This provided molecule-scale spatial localization and information on molecular orientation. The potential biological applications of single-molecule imaging captured the imaginations of microscopists and biologists alike, but because of its invasiveness and complexity NSOM proved largely unsuitable for complex biological samples.
http://www.nature.com/milestones/milelight/full/milelight17.html
"Nearfield optics" by Zogdog602 - Own work. Licensed under CC BY 3.0 via Wikimedia Commons - http://commons.wikimedia.org/wiki/File:Nearfield_optics.png#mediaviewer/File:Nearfield_optics.png

## Slide 52: NSOM working in the near-field
Near-field near surface of object, < λ of light
Near-field consists of light as evanescent wave
Evanescent waves higher frequency, more information
Evanescent waves quantum tunneling phenomenon
Product of Schrödinger wave equations
Notes: http://micro.magnet.fsu.edu/primer/techniques/nearfield/nearfieldintro.html
http://www.azonano.com/article.aspx?ArticleID=1205

## Slide 53: Near-Field Scanning Optical Microscopy (NSOM) How to make an NSOM tip
Tip of pulled quartz fiber
Very small fraction of light makes it through small (50nm) aperture
Aluminize tip to minimize loss of light

## Slide 54: Near-Field Scanning Optical Microscopy (NSOM)
SEM of tip
Tip shining on sample
(can detect with wide-field)

## Slide 55: How to move the tip? Steal from AFM
Atomic Force Microscopy (AFM)
Notes: Sharp tip at end of swinging cantilever
http://simple.wikipedia.org/wiki/Atomic_force_microscope

## Slide 56: Near-Field Scanning Optical Microscopy (NSOM) Break the diffraction limit by working in the near-field
Like AFM can do NSOM with tapping mode
Requires bent tip
Move tip up and down like AFM
Not best way of doing NSOM
  Hard to make probe
  Bend causes loss of light
Notes: There are several drawbacks in the application of bent optical probes, each of which can be attributed to the bend itself. A significant problem is the increased difficulty of the probe fabrication, especially when applying a metal coating to the tip. An additional disadvantage is the increased optical loss that occurs due to the bend in the probe. This loss in throughput efficiency is significant, and some published measurements indicate that bent optical fiber probes are at least an order of magnitude less efficient than conventional straight fiber probes. In certain operational modes of NSOM, the intensity loss is not a serious limitation because additional light can be coupled into the fiber to compensate, assuming sufficient laser power is available. The increase in optical coupling is an option because the optical losses, as well as increased heating, occur at the bend in the fiber and not at the aperture of the probe, where local heating would present a major problem. Another potential drawback with bent probes is a change in certain tip properties that occurs due to the presence of the bend, such as a decrease in extinction coefficients when performing polarized light measurements. Extinction ratios of approximately 70:1 have been measured in the far-field utilizing bent tips, as compared to values of greater than 100:1 with conventional straight fiber probes.

## Slide 57: If not tapping like AFM how else to scan tip in NSOM? Shear force mode. Advantage: don’t need laser to keep track of probe.
To keep tip in near-field, need to be ~50nm from surface
Notes: The main problem associated with this type of feedback mechanism is that the light source (for example, a laser), which is used to detect the tip vibration frequency, phase, and amplitude, becomes a potential source of stray photons that can interfere with the detection of the NSOM signal. One mechanism for dealing with this effective increase in background signal is to provide a feedback light source that has a different wavelength (usually longer) than the near-field source. This scheme requires additional filtration in front of the detector to selectively block the unwanted photons originating within the feedback system. In most cases, the added filters also block a small percentage of the near-field photons, resulting in reduced signal levels. A non-optical feedback method is not subject to problems of this nature, and is a primary reason that methods such as the tuning-fork technique (described below) have become increasingly popular.

## Slide 58: Sense presence of surface from dithering tip (lateral) (Increased shear force when surface is near)
Keep dithering amplitude low <10 nm
Notes: Betzig used this approach in 1992 to make NSOM more useful
The shear-force feedback method laterally dithers the probe tip at a mechanical resonance frequency in proximity to the specimen surface. The dither amplitude is usually kept low (less than 10 nanometers) to prevent adversely affecting the optical resolution. For optimum image quality, shear-force feedback techniques are usually restricted to use with specimens that have relatively low surface relief, and longer scan times are required compared to operation in tapping mode. However, the straight probes typically employed in shear-force feedback techniques are easier to fabricate and have a lower cost per probe than their bent probe counterparts.
http://micro.magnet.fsu.edu/primer/techniques/nearfield/nearfieldintro.html
http://zeiss-campus.magnet.fsu.edu/articles/superresolution/introduction.html

## Slide 59: Shear force mode with non optical feedback
Use real-time feedback to keep probe in near-field range but not touching
Tip can be oscillated at resonance frequency
Tip can be straight
  Easier to make
  Cheaper
  But surface needs to be relatively flat
Notes: There are several advantages of the tuning fork method that have led to its increased favor over optical techniques of tip regulation. Since the detection of the tip motion is not optical, there is no risk of additional stray light being introduced in the vicinity of the aperture that might interfere with the NSOM signal detection. Additionally, the tuning fork system does not require the tedious alignment procedures of a separate external laser source and associated focusing optical components. Because of the compactness and relative ease of use, the tuning fork method lends itself to applications requiring remote operation, such as those employed in vacuum systems or environmental control chambers.
The basic configuration of the tuning-fork method used for shear-force tip feedback consists of a single mode optical fiber attached to one arm of a quartz crystal tuning fork, which is oscillated at the tuning fork's resonance frequency. The equivalent circuit for the tuning fork is a series RLC resonator in parallel with package capacitance. The most common tuning fork resonance frequency is 32,768 hertz (Hz), but the devices are available with resonances ranging from 10 kilohertz to several tens of megahertz.

## Slide 60: Illumination Techniques - Overview
Transmitted Light
Bright-field
Oblique
Darkfield
Phase Contrast
Polarized Light
DIC (Differential Interference Contrast)
Fluorescence - not any more > Epi !
Reflected (Incident) Light
Bright-field
Oblique
Darkfield
Not any more (DIC !)
Polarized Light
DIC (Differential Interference Contrast)
Fluorescence (Epi)

## Slide 61: NSOM, like far-field, is amenable to different contrast methods
Absorption
Polarization
Refractive index
Fluorescence
Spectral imaging
Reflected
Reflected Light
Transmitted Light
Notes: Betzig, E., Trautman, J.K., 1992. Near-Field Optics: Microscopy, Spectroscopy, and Surface Modification Beyond the Diffraction Limit. Science 257, 189-195.

## Slide 62: Direct imaging of single molecule with NSOM (1993)
Instrument described in 1992 Science paper
Shear force mode with non-optical feedback
In 1993 Eric Betzig and Robert Chichester used NSOM for repetitive single molecule imaging, DiI
Notes: carbocyanine dye
In 1993, this situation changed markedly when Eric Betzig and Robert Chichester reported the first repetitive imaging of single fluorophores at room temperature with a new technique called near-field scanning optical microscopy (NSOM) that repeatedly scans an extremely small optical probe over a sample. This provided molecule-scale spatial localization and information on molecular orientation. The potential biological applications of single-molecule imaging captured the imaginations of microscopists and biologists alike, but because of its invasiveness and complexity NSOM proved largely unsuitable for complex biological samples.
http://www.sciencemag.org/content/257/5067/189.abstract?sid=78a939a7-a92c-4260-a999-85346db22447
Betzig, E., Trautman, J.K., 1992. Near-Field Optics: Microscopy, Spectroscopy, and Surface Modification Beyond the Diffraction Limit. Science 257, 189-195.
http://www.sciencemag.org/content/262/5138/1422.short
Betzig, E., Chichester, R.J., 1993. Single Molecules Observed by Near-Field Scanning Optical Microscopy. Science 262, 1422-1425.
http://www.nature.com/milestones/milelight/full/milelight17.html
Fig. 1. Near-field scanning optic
microscopy : conception and real­ ity. (A) As originally conceived ( 10), an illuminated aperture acts as a light source of subwavelength dimensions that can be scanned close to an object to generate a superresolution image. (B) As we implemented it, the aperture is at the end of a sharp probe,and shear forces between the tip and sample
  are measured to automatically control the relative separation for scanning rough surfaces . The scan
head is designed as an attachment to a commercial optical microscope, which results in a wide range of possible magnifications . The near-field signal can also be partitioned to obtain information from several contrast mechanisms at once.

## Slide 63: NSOM images Single molecules of DiI on glass surface

## Slide 64: NSOM images

## Slide 65: NSOM disadvantages

## Slide 66: NSOM disadvantages
Practically zero working distance and small depth of field.
Extremely long scan times for high resolution images or large specimen areas.
Very low little light through such a tiny aperture.
Only features at surface of specimens can be studied.
Fiber optic probes are somewhat problematic for imaging soft materials due to their high spring constants, especially in shear-force mode.

## Slide 67: Performance range of optical microscopy
CLSM
Depth
(um)
Resolution
(um)
LM
OCT
NSOM
MRI
SPIM
SIM/STP
TIRF
Notes: serial two-photon (STP) tomography
http://www.nature.com/nature/journal/v508/n7495/full/nature13186.html#methods
http://www.nature.com/nmeth/journal/v9/n3/full/nmeth.1854.html

## Slide 68: Critiquing figures

## Slide 69: Critiquing figures

## Slide 70: Critiquing figures