## Slide 1: Biology 177: Principles of Modern Microscopy
Lecture 06:
Fluorescence Microscopy

## Slide 2: Lecture 6: Fluorescence Microscopy
Detectors for Microscopy, Part 2
  CMOS, PMT and APD
Phenomenon of Fluorescence
  Energy Diagram
  Rates of excitation, emission, ISC
  Practical Issues
    Lighting, Filters
Homework 2 review

## Slide 3: Detectors for microscopy
Film
CMOS (Complementary metal–oxide–semiconductor)
CCD (Charge coupled device)
PMT (Photomultiplier tube)
GaAsP (Gallium arsenide phosphide)
APD (Avalanche photodiode)
Array of detectors, like your retina
Single point source detectors

## Slide 4: Let’s look at an actual example
Visualizing hearing in vivo

## Slide 5: High speed cameras
Used by the military so expensive
10,000 frames/sec.
Fastcam 1024 PCI
  Photron, up to 100,000 fps
  47,000 with LED illumination and DIC
ProAnalyst software
  Xcetex, Cambridge, Mass
Actually Fastcam SA5

## Slide 6: Nyquist criterion
Sampling frequency needs to be greater than twice the frequency trying to image
Undersampling can result in aliasing
Such differences can result in distortions or artifacts
  Avoid the “wagon wheel” effect
Notes: In terms of a function's own bandwidth (B), as depicted above, the Nyquist criterion is often stated as fs > 2B. And 2B is called the Nyquist rate for functions with bandwidth B. When the Nyquist criterion is not met (B > ½ fs), a condition called aliasing occurs, which results in some inevitable differences between x(t) and a reconstructed function that has less bandwidth. In most cases, the differences are viewed as distortion.

## Slide 7: Visualizing Hearing In Vivo
Spontaneous otoacoustic emissions
  In vivo SOAE 600 Hz to 60 kHz
  In vitro hair bundle motions < 100 Hz
Playing sounds 100 to 2000 Hz
Experimental Setup
  High magnification DIC microscopy
  High speed camera 6000 fps +
  Specialized data analysis software
Glass bottom dish
Speaker
Specimen
Water-filled tube

## Slide 9: Temporal Resolution: Ultra High-Speed Video
1000 fps is 1024x1024
6000 fps reduced to 512x256, 1024x128, etc.
High Power LED System-36AD3500, Lightspeed Technologies, Campbell, CA
  20/80 to 90/10 on/off
Fastcam SA1 (Photron, San Diego, CA), 5400 fps at 1024x1024

## Slide 10: 6000 fps movie played at 30 fps

## Slide 11: Med Engineering Seminar Last Week
Lihong V. Wang
Working on Detector that can go 1,000,000,000 frames/sec!
So we all know the problem here

## Slide 12: Med Engineering Seminar Last Week
Lihong V. Wang
Working on Detector that can go 1,000,000,000 frames/sec!
So we all know the problem here

## Slide 13: Photomultiplier Tube (PMT)
Two types of Photomultipliers
  Side-on, most popular due to high performance rating and low cost
  Head-on
Notes: Point source detector
The two most popular photomultiplier designs have configurations that place the sensitive photocathode element either at the end (head-on) or on the side (side-on) of the vacuum tube (see Figure 3). Incident light is detected through a window at the top of the glass envelope in head-on designs, and through the curved side in side-on photomultipliers. Due to their high performance ratings and low cost, side-on photomultipliers are the most widely used tubes for general photometric applications, such as spectrophotometry, fluorimetry, and confocal microscopy. Side-on photomultipliers contain an opaque and relatively thick reflection-mode photocathode surrounded by a circular cage-like dynode element chain. Incident photoelectrons do not pass through the photocathode in a side-on photomultiplier, but instead are ejected from the front face and angled toward the first dynode element.
In contrast, the photocathode in an end-on photomultiplier must be of precise thickness as well as composition, and is semi-transparent. The photoemissive material is deposited on the surface of an optical window so that electrons are emitted from the side of the photocathode opposite that of the incident radiation. If the photocathode is too thick, more photons will be absorbed but fewer photoelectrons will be emitted from the rear surface. Alternatively, if the photocathode is too thin, photons may pass straight through without being absorbed. Because of these differences in design, side-on photomultipliers having photocathodes of similar composition to head-on tubes often exhibit higher quantum efficiencies.

## Slide 14: Photomultiplier Tube (PMT)
PMTs use electric potential to amplify electrons
Photons impact a phosphor screen creating electrons
Electrons are multiplied by impacting other surfaces (Dynode chain)
Increasing the gain increases the number of electrons produced in a non-linear fashion
So increasing Gain increases signal

## Slide 15: Photomultiplier Tube (PMT)
PMTs less sensitive in the red
Can buy PMTs that are more “green” sensitive or more “red”sensitive
  But not much difference
Notes: Even with their superior gain characteristics, photomultipliers have some disadvantages when compared to other types of detectors. The gain and dark current performance of individual photomultipliers from the same production run can vary widely, often by a factor of 2 to 5 for a particular design. In addition, the devices can be damaged by exposure to high illumination levels, often requiring very long recovery times (some never recover from high intensity illumination). Modern photocathodes yield very good visible and ultraviolet performance, but sensitivity drops dramatically in the longer visible wavelengths and infrared. Photomultipliers require a stabilized high voltage power supply, which increases the cost of using these detectors. Finally, even in the most sensitive wavelength region, about two-thirds of the incident light is not detected by the photocathode and does not contribute to photomultiplier gain.

## Slide 16: As with CCDs, PMTs have same Noise problems
Shot noise
  Random fluctuations in the photon population
Dark current
  Noise caused by spontaneous electron formation/accumulation in the wells (usually due to heat)
Readout noise
  Grainy noise you see when you expose the chip with no light

## Slide 17: The cost of increasing gain
More electrons means more noise
This is what causes the noise in scanning confocal images
  Averaging can decrease the noise
Notes: http://microscopy.duke.edu/learn/introtomicroscopy/confocals.html

## Slide 18: Dynamic Range (bit depth)
Full well capacity/read noise
2^8 = 256 gray values = 8 bits, 2^10 = 1024 grays values = 10 bits, etc
8 bit (video camera)
  Your eye can detect this range, computers and printers are therefore designed around this value
16 bit
  Good to detect very dim and very bright things in the same field of view without saturation (maxing out the range)
  Very good for quantifying fluorescence
  Have to convert to 8 bits for presentations
Notes: Applies to PMT and CCD

## Slide 19: Confocal Imaging: Avalanche Photodiodes
Cindy Chiu from David Prober’s Lab, Caltech
Notes: APD not linear. This can be good.

## Slide 20: Fluorescence Microscopy
The Ultimate in Contrast

## Slide 21: Thus Far, have considered compound microscope, and the microscope optics as a projection system (into eye)
Deliver light to the specimen
Image light from the specimen
Contrast from light absorbed, diffracted

## Slide 22
Transmitted light microscopy: photons out of the microscope are some fraction of the photons in
Now, turn our attention to fluorescence, based on the absorption and re-emission of photons
Fluorescent Dye
Dipole antenna
Delocalized electrons
Longer dipole, longer l

## Slide 23: Fluorescence
Easy to set up: Objective = Condenser
Highly specific technique, wide selection of markers
Detection and Identification of Proteins, Bacteria, Viruses
Basics for
  Special Techniques eg. TIRF, FRET, FRAP etc.
  3-D imaging
  Deconvolution
  Structured Illumination
  Confocal Techniques

## Slide 24: Light sources
Mercury (Hg)
Xenon, Hg/Xe Combination
Laser
LED’s
Tungsten Halogen

## Slide 25
Dye in cuvette
Blue light absorbed
Beer’s Law
Iout = Iin e-ax
Iabsorbed = Iout - Iin
= Iin(1-e-ecx)
e = extinction coefficient
For Fluorescein
e ~ 70,000/(cm M/liter)
490nm
Light absorbed
Wavelength
A good dye must absorb light well (high extinction coef.)
Notes: Beer Lambert Law
x = distance light travels through the material
α = attenuation coefficient
The attenuation coefficient is a quantity that characterizes how easily a material or medium can be penetrated by a beam of light. A large attenuation coefficient means that the beam is quickly "attenuated" (weakened) as it passes through the medium, and a small attenuation coefficient means that the medium is relatively transparent to the beam.
α = ε c
c = concentration of attenuating species in material
ε =extinction coefficient also called molar absorption coefficient, absorptivity of attenuator

## Slide 26
Blue light absorbed
490nm
520nm
Green light emitted
Stokes Shift
Where does energy go?
Quantum Yield = light emitted/light absorbed
Q ~ 0.8 fluorescein
~ 0.3 rhodamine

## Slide 27
Co-fluorescein
Co-TM rhodamine

## Slide 28
Which dye is better?
1 - absorb well (high e )
2 - emit well (high Q)
Brightness ~ eQ
(fluorescein 0.8 * 70,000 = 57,000)
(rhodamine 0.3 * 90,000 = 27,000)

## Slide 29
4nsec
Go deeper to explain bleaching and background
(Jablonski diagram)
0.8 emitted
Other losses
Heat
Energy transfer

## Slide 30
4nsec
Add in Interstate Crossing (ISC)
0.8 emitted
fluorescence
ISC
~0.03
Excited triplet
state
Phosphorescence
(usec - msec)
Triplet state is long lived.
therefore even low probability can deplete active dye
(steady state reached in ~200msec
~80-90% in triplet --> 5-10 fold dimmer)
CLSM: can have a major impact (~5 fold less throughput)

## Slide 31
4nsec
Interstate Crossing (ISC) Problem 2: Reactive oxygen
0.8 emitted
fluorescence
ISC
~0.03
Excited triplet
state
Phosphorescence
(usec - msec)
Triplet state lifetime shortened by oxygen
(20msec if none; 0.1 usec if oxygen present
Good news: Returns dye to ground state
Bad news: Creates reactive oxygen

## Slide 32
Aside: Phosphor Imager
ISC
High probability
Excited triplet
state
Phosphorescence
(very slow)
Very
slow
Accumulate triplet state (thermally stable)
“read out” with scanning red laser
Gives energy for transition to singlet state
Emission of light proportional to the stored triplet

## Slide 33
Issues in fluorescence
1. No dye is perfect < 100,000 photons total
(ISC, bleaching)
2. Every emitted photon is sacred
(NA 1.25 collects ~20%)
(clsm w/ PMT collects 0.02% - 0.3%)
3. Signal/noise limited by number of photons
Counting error N ± sqrt(N)
Image requires >200 photons/pixel

## Slide 34
Not enough fluorescence photons?
If >200 photons/ pixel needed
Microscope records 0.02%
Need about 100,000 photons/pixel
~ lifetime of a dye
Given dwell-time of laser beam, ISC, collection efficiency
Lucky to record 1 photon/dye/scan
Every emitted photon is sacred!
Maximize throughput (filters, lenses, mirrors)
Minimize Bleaching

## Slide 35
To reduce bleaching:
Shorten Triplet lifetime
Antibleach Agents:
Retinoids, carotinoids, glutathione
Vitamin E, N-propyl gallate
Eliminate Oxygen (scavenger, bubble N2)
No reactive oxygen produced
(but lengthens triplet lifetime)

## Slide 36
Can’t get more light by turning up the laser:
Dye saturates as I is increased
Intense laser beam depletes dye in ground state
Pumps more dye into the triplet state
(reactive oxygen and silent)
Noise doesn’t saturate
Autofluorescence in cell
flavins, NADH, NADPH
Raman spectrum of water
(488nm in; 584nm out)

## Slide 37
Optimize light collection, uniformity of illumination
High NA, Kohler illumination

## Slide 38
N.A. and image brightness
Transmitted light
Brightness = fn (NA2 / magnification2)
Epifluorescence
Brightness = fn (NA4 / magnification2)
10x 0.5 NA is 3 times brighter than 10x 0.3NA
10x 0.5 NA is 8 times brighter than 10x 0.3NA
q
N.A. = h sin q

## Slide 39: First Fluorescence microscope
Built by Henry Seidentopf & August Köhler (1908)
Used transmitted light path
So dangerous that couldn’t look through it, needed camera
Image credit: corporate.zeiss.com “Technical Milestones of Microscopy”
Notes: Inspired by Ernst Abbe's finding that shorter wavelength light leads to higher resolution, August Köhler constructed the first ultraviolet (UV) microscope at Zeiss Optical Works in Jena, Germany, in 1904. The instrument used UV illumination from a cadmium arc lamp and allowed photographic reproduction of objects at twice the resolution of visible light microscopes. Köhler also noted that some objects emitted light of longer wavelength on illumination with UV, but it was the physicist Oskar Heimstädt who, in 1911, used this observation as the basis for the construction of the first successful fluorescence microscope.

## Slide 40: First epi-fluorescence microscope
Designed in 1929 by German pharmacologist Philipp Ellinger & anatomist August Hirt
Used yellow barrier filter between objective and ocular to block reflected excitation light
Would be custom made, specialized instrument for almost 40 years
Notes: In what they called an 'intravital microscope' the excitation light passed through a series of filters before the right wavelength hit the objective lens and triggered emission of fluorescent light in the observed tissue. A yellow barrier filter between the objective and ocular prevented most reflected excitation light from interfering with observation. Ellinger and Hirt imaged kidney and liver tissue in rodents injected with the fluorochromes fluorescein and trypaflavin.
http://www.nature.com/milestones/milelight/full/milelight04.html

## Slide 41: Key advance dichroic mirrors!
Dutch scientist Bas Ploem developed these in 1967
Dichromatic mirrors converted epi-fluorescence microscope from a tool that could be used only by trained specialists to a universal and indispensable instrument for modern biology
Prototype of first epi-illumination fluorescence microscope that he developed in Amsterdam in the sixties
Notes: http://www.cyto.purdue.edu/cdroms/cyto10a/cytometryhistory/individualhistories/tanke.html
http://www.cyto.purdue.edu/cdroms/cyto10a/cytometryhistory/individualhistories/ploem.html
http://www.nature.com/milestones/milelight/full/milelight04.html

## Slide 42
Choose filters well
Excitation
Dichroic
Emission
Optimize the light path for collection

## Slide 43
Excitation filter:
Selectively excite dye
Emission filter:
Selectively detect dye
Dichroic Reflector:
Bounce exciting l
Pass emitted l

## Slide 44
How to separate wavelengths: Interference Filters
Basic principle based on reflection from mirror
Reflection from higher index --> 180 degree shift
(separated for clarity below)
mirror
Notes: A dichroic filter, thin-film filter, or interference filter is a very accurate color filter used to selectively pass light of a small range of colors while reflecting other colors. By comparison, dichroic mirrors and dichroic reflectors tend to be characterized by the color(s) of light that they reflect, rather than the color(s) they pass. (See dichroism for the etymology of the term.)

## Slide 45
Interference Filters
Add a layer of intermediate index
3% reflection from glass
(higher index --> 180 degree shift
(separated for clarity below)
Constructive interference
Less light passed
2
Note: thickness of layer in terms of wavelength
Notes: A dichroic filter, thin-film filter, or interference filter is a very accurate color filter used to selectively pass light of a small range of colors while reflecting other colors. By comparison, dichroic mirrors and dichroic reflectors tend to be characterized by the color(s) of light that they reflect, rather than the color(s) they pass. (See dichroism for the etymology of the term.)

## Slide 46
Interference Filters are wavelength dependent
 2 = 2 x  1
Destructive interference
(antireflection coating)
most light passed
4
Same thickness is smaller in terms of wavelength for  2
 2
 1
Constructive interference
Less light passed
2
Notes: A dichroic filter, thin-film filter, or interference filter is a very accurate color filter used to selectively pass light of a small range of colors while reflecting other colors. By comparison, dichroic mirrors and dichroic reflectors tend to be characterized by the color(s) of light that they reflect, rather than the color(s) they pass. (See dichroism for the etymology of the term.)

## Slide 47: Interference filters: the movie
Reflect one wavelength while passing another
Innovation that relatively inexpensive to make
Link to Java Tutorial
http://www.olympusfluoview.com/theory/interferencefilters.html
Notes: Recent technological achievements in bandpass filter design have led to the relatively inexpensive construction of thin-film interference filters featuring major improvements in wavelength selection and transmission performance. These filters operate by transmitting a selected wavelength region with high efficiency while rejecting, through reflection and destructive interference, all other wavelengths.

## Slide 48
Dichroic reflector
Issues:
How steep,
How efficient to excite
How efficient to collect
Dichroic reflectors tend to be characterized by the color(s) of light that they reflect, rather than the color(s) they pass
Notes: Also called Dichroic mirror

## Slide 49
Excitation filter:
Selectively excite dye
Emission filter:
Selectively detect dye
Dichroic Reflector:
Bounce exciting l
Pass emitted l

## Slide 50: Epi - Fluorescence
(Specimen containing green fluorescing Fluorochrome)
Dichromatic Mirror
Emission Filter
Excitation Filter
Observation port
FL
Light Source
Specimen containing green fluorescing Fluorochrome

## Slide 51: Here is what they look like
Nikon
Olympus

## Slide 52: Different kinds of Emission and Excitation Filters

## Slide 53: Reading bandpass filter spectrum
All have a center wavelength
Guaranteed Minimum Bandwidth (GMBW)
This is less than the FWHM
Example 520/35 filter (502.5-537.5)
Notes: http://www.semrock.com/bandwidth.aspx
Full Width at Half Maximum (FWHM)

## Slide 54
Final Note:
Resonance Energy Transfer (non-radiative)
The Bad: Self-quenching
If dye at high concentration
“hot-potato” the energy until lost

## Slide 55
Final Note:
Resonance Energy Transfer (non-radiative)
The Good: FRET as a molecular yardstick
Transfer of energy from one dye to another
Depends on:
Spectral overlap
Distance
Alignment

## Slide 56
FRET:
Optimize spectral overlap
Optimize k2 -- alignment of dipoles
Minimize direct excitement of the acceptor
(extra challenge for filter design)
donor
acceptor

## Slide 57: Homework 2
The answer.

## Slide 58: The Finitely Corrected Compound Microscope
Objective
Eyepiece
Objective Mount (Flange)
150 mm (tube length = 160mm)
B
B
A
In most finitely corrected systems, the eyepiece has to correct for the Lateral Chromatic Aberrations of the objectives, since the intermediate image is not fully corrected.
(Note: the LCA correction is done in a brand-specific fashion)

## Slide 59
Tube lens
(Zeiss: f=164.5mm)
Objective
Eyepiece
The Compound Microscope (infinity corrected)

## Slide 60
Homework 2: Why are most modern microscopes “infinity corrected”
Hint - think of the influence of a piece of glass
Image
Eyepiece
image
Eyepiece
Lens of eye

## Slide 61
Image
Eyepiece
image
Eyepiece
Lens of eye
Take special case:
Glass at right angle to second principle ray
Simplify by removing eyepiece and eye

## Slide 62
Image
Eyepiece
image
Take special case:
Glass at right angle to second principle ray
Refraction of principle rays
Zone of Confusion: Rays fail to intersect at only one place

## Slide 63
Image
Eyepiece
image
Eyepiece
Tube lens
Objective
“Infinity” Domain
“Infinity correction” provides a region in which an optical flat will not create a zone of confusion
Lens of eye

## Slide 64
Infinity optics creates a domain in which all rays from same point in object are parallel
Infinity domain
Good Aspects:
Optical flats inserted have no effect (shift doesn’t matter)
Magnification unchanged by adding accessories
BUT:
Remember that thin lens laws no longer apply
http://microscopy.fsu.edu/primer/anatomy/infinityintro.html

## Slide 65
Nikon. Leica
Zeiss
Different manufacturers have elected different compromises
  Length of objective lens
  Diameter of objective lens
  Focal length of tube lens
Longer tube lens focal length easier to design,
But requires larger diameter threads.

## Slide 66
Conjugate Planes in Infinity Optics
Illumination Path
Imaging Path
Eyepiece
TubeLens
Objective
Condenser
Collector
Eye
Field Diaphragm
Specimen
Intermediate Image
Retina
Light Source
Condenser Aperture Diaphragm
Objective Back Focal Plane
Eyepoint

## Slide 67: Fluorescent proteins
Proteins from marine invertebrates
Can be coded in genes and made by the organism
Now come in a variety of colors

## Slide 68: Green Fluorescent Protein
First fluorescent protein discovered and developed for biological use
Mutated for temp stability, color and turnover rate
Importance of monomer vs dimer or tetramer
Notes: http://www.olympusconfocal.com/applications/fpcolorpalette.html
http://www.micro.magnet.fsu.edu/primer/techniques/fluorescence/fluorescentproteins/fluorescentproteinshome.html

## Slide 69: Photoconvertible Proteins
Kaede, coral fluorescent protein, tetramer
Dendra2, from soft coral, monomer
UV Laser (405 nm) to convert green to red
ROI (Region Of Interest) allows precise targeting
www.olympusfluoview.com
www.amalgaam.co.jp