## Slide 1: Biology 177: Principles of Modern Microscopy
Lecture 12:
Fluorescent labeling, multi-spectral imaging and FRET

## Slide 2: Lecture 12: Fluorescent labeling, multi-sprectral imaging and FRET
How to characterize the performance of fluorescent probes?
“Quantitative“ Fluorescence
  Fluorescence linearity (non-linearity)
    Dye, microscope, camera
  Flat-fielding to linearize
  Quantitating the image
  Multispectral imaging
  FRET

## Slide 3
Our discussion of fluorescence has made hidden assumption that dyes have an ideal behavior
How true is this?
Fluorescent Dye
Dipole antenna
Delocalized electrons
Longer dipole, longer l
Notes: http://micro.magnet.fsu.edu/primer/techniques/fluorescence/fluorescenceintro.html
http://olympusmicro.com/primer/java/jablonski/solventeffects/index.html

## Slide 4
Dye in cuvette
Blue light absorbed
Beer’s Law
Iout = Iin e-ax
Iabsorbed = Iout - Iin
= Iin(1-e-ecx)
e = extinction coefficient
For Fluorescein
e ~ 70,000/(cm M/liter)
490nm
Light absorbed
Wavelength
A good dye must absorb light well (high extinction coef.)
Notes: Beer Lambert Law
x = distance light travels through the material
α = attenuation coefficient
The attenuation coefficient is a quantity that characterizes how easily a material or medium can be penetrated by a beam of light. A large attenuation coefficient means that the beam is quickly "attenuated" (weakened) as it passes through the medium, and a small attenuation coefficient means that the medium is relatively transparent to the beam.
α = ε c
c = concentration of attenuating species in material
ε =extinction coefficient also called molar absorption coefficient, absorptivity of attenuator

## Slide 5: Fluorophore absorption
Beer-Lambert law
Iout = Iin exp (-  L c)
Iabsorbed = Iout - Iin
Iin: incident light intensity (in W.cm-2)
L: absorption path length (in cm)
c: concentration of the absorber (in M or mol.L-1)
: molar absorption coefficient (in M-1cm-1 or mol-1.L.cm-1)
Fluorescein  ~ 70,000 M-1.cm-1
eGFP  ~ 55,000 M-1.cm-1
Light absorbed
Wavelength
Iin
Blue light absorbed
Iout
Green dye in cuvette
L
Notes: Classically, the Beer–Lambert law was first devised independently where Lambert's law stated that absorbance is directly proportional to the thickness of the sample, and Beer's law stated that absorbance is proportional to the concentration of the sample. The modern derivation of the Beer–Lambert law combines the two laws and correlate the absorbance to both, the concentration as well as the thickness (path length) of the sample.

## Slide 6: Fluorophore absorption
Other expressions of the Beer-Lambert law:
Iout = Iin exp (-  L c)
Iout = Iin exp (-  L N)
Iout = Iin exp (- µa L)
Iin: incident light intensity in W.cm-2
L: absorption path length in cm
c: concentration of the absorber in M or mol.L-1
N: density of the absorber in molecule.cm-3
: molar absorption coefficient in M-1cm-1 or mol-1.L.cm-1
: absorption cross section in cm2 or cm2.molecule-1
µa: absorption (attenuation) coefficient in cm-1
N = NAvogadro . 10-3 c (1L = 103 cm3)
 = NAvogadro . 10-3  = 6.022 1020 
eGFP
 = 55,000 M-1.cm-1
 = 9.13 10-17 cm2.molecule-1
Notes: To describe the attenuation coefficient in a way independent of the material density, one introduces the cross section σ = μ/n, where n is the numerical density (number of atoms per volume) of the material. σ has the dimension of an area; it expresses the likelihood of interaction between particles.
The attenuation (or absorption) coefficient is a quantity that characterizes how easily a material or medium can be penetrated by a beam of light. A large attenuation coefficient means that the beam is quickly "attenuated" (weakened) as it passes through the medium, and a small attenuation coefficient means that the medium is relatively transparent to the beam.

## Slide 7: Fluorophore absorption
µt = µa + µs
µt= c
µa = a c
 = a if no scattering
absorption
coefficient
scattering
coefficient
extinction
coefficient
molar
absorption
coefficient
molar
extinction
coefficient
In the literature…
The “extinction coefficient” is usually given in tables.
confusions:
- “extinction coefficient” used for “absorption coefficient” (it assumes the scattering coefficient is negligible)
- “extinction coefficient” used for “molar extinction coefficient” (check the unit!)
()!
The maximum is given in tables, or the excitation wavelength is indicated.

## Slide 8: Fluorophore absorption
it is the molar absorption coefficient
Shaner et al, Nature Biotechnology, 2004
Example: Properties of fluorescent protein variants

## Slide 9: Fluorophore absorption
Light absorbed
Wavelength
Iin
Blue light absorbed
Iout
Green dye in cuvette
L
Green light emitted
Iemitted
490nm
520nm
Stokes Shift
Light emitted
Quantum Yield
Q = Iemitted /Iabsorbed
= # photons emitted / # photons absorbed
(Iabsorbed = Iout - Iin)
Fluorescein Q ~ 0.8
Rhodamine B Q ~ 0.3
eGPP Q ~ 0.6

## Slide 10: Fluorophore brightness = Q
Example: Properties of fluorescent protein variants
DsRed Q ~ 0.79 x 75,000 ~ 59,250 M-1.cm-1 (100%)
mRFP1 Q ~ 0.25 x 50,000 ~ 12,500 M-1.cm-1 (21%)
eGFP Q ~ 0.6 x 55,000 ~ 33,000 M-1.cm-1 (56%)
Fluorescein Q ~ 0.8 x 70,000 ~ 56,000 M-1.cm-1 (95%) (dye!)
Notes: Shaner et al, Nature Biotechnology, 2004

## Slide 11: The dilute limit
Extinction coefficient and quantum yield corresponds to “well behaved” dye in the dilute limit: dilute photon and dilute dye
From Michael Liebling, UCSB

## Slide 12: The dilute limit: dilute photons
As photons hit specimen: dye molecules excited and less dye left unexcited
Dye remaining to be excited
Excitation intensity
Emission Intensity
Excitation intensity
Saturation!

## Slide 13: Interstate (or Intersystem) crossing (ISC)
From Michael Liebling, UCSB

## Slide 14: Interstate crossing (ISC) and photobleaching
From Michael Liebling, UCSB
As ISC takes place: less dye molecules available and unexcitable dye accumulates

## Slide 15: Cycle of a fluorophore
From Michael Liebling, UCSB

## Slide 16: Interstate crossing (ISC) and photobleaching
A good dye is more photostable (less photobleaching)

## Slide 17: Interstate crossing (ISC) and photobleaching
after 30 seconds
Bovine pulmonary artery endothelial cells (BPAEC) were labeled with fluorescein phalloidin (left panels, Cat. no. F432), or Alexa Fluor® 488 phalloidin (right panels, Cat. no. A12379), which labels filamentous actin, and mounted in PBS. The cells were placed under constant illumination on the microscope with an FITC filter set using a 60× objective. Images were acquired at one-second intervals for 30 seconds. Under these illumination conditions, fluoresce in photobleached to about 20% of its initial value in 30 seconds; the fluorescence of Alexa Fluor® 488 phalloidin stayed at the initial value under the same illumination conditions.
Fluorescein
Alexa 488
Notes: http://www.lifetechnologies.com/us/en/home/technical-resources/research-tools/image-gallery/image-gallery-detail.altid.g001270.html

## Slide 18: Photobleaching characterization
Shaner et al, Nature Biotechnology, 2004
Example: Properties of fluorescent protein variants
Notes: Bleaching in seconds; obviously longer better, to a point

## Slide 19
Resonance Energy Transfer (non-radiative)
Transfer of energy from one dye to another
Depends on:
Spectral overlap
Distance
Alignment
Notes: http://www.olympusmicro.com/primer/techniques/fluorescence/fret/fretintro.html

## Slide 20
RET is not always between dissimilar dyes
“Self-quenching” of dye
(“hot-potato” the energy until lost)
Depends on:
Dye Concentration
Geometry
Environment
Log [dye]
Log I
~0.1uM
Notes: In a living biological sample can be difficult to reach such high dye concentration without killing cells or making them sick.

## Slide 21
“Self-quenching” of dye
Depends on:
Dye Concentration
Geometry
Log [dye]
Log I
~0.1uM
Hard for emission from this one
Easier emission from this one
A uniformly dyed structure
Instead, looks “hollow”

## Slide 22: Fluorescence quantification based on signal intensity
Example:
in = level of expression of a fluorescent protein
out = fluorescent signal and grey level of pixel on an image.
input: [fluorophore]
Output: pixel grey levels
Notes: Want higher bit depth. Do not bother with 8 bit images even though using it in these examples.

## Slide 23: Example of nonlinearity: Pixel saturation
(detector or digital contrast)
Solution(s): use less power!!!, decrease the acquisition time, decrease [fluorophore],…
Notes: What is the bit depth?

## Slide 24: Example of nonlinearity: Fluorophore saturation
Solution(s): use less power!!!, decrease the acquisition time,…

## Slide 25: Example of nonlinearity: Noise
Solution(s): optimize the excitation wavelength, increase the acquisition time, use more power, use a stronger fluorophore, increase [fluorophore],…

## Slide 26: Example of nonlinearity: Photo-induced fluorescence
Solution(s): use less power!

## Slide 27: Example of nonlinearity: Photo-induced fluorescence
Red channel
zoom out after imaging in this area
Green channel
Bleaching!
Induced
auto-fluorescence!

## Slide 28: Compare what is comparable: Imaging depth
Same object imaged at different tissue depth…
The fluorescence level depends on
the depth of imaging and the optical properties of the tissue
(variation from one sample to another)…
All dyes look redder as you look deeper in tissue

## Slide 29: How to protect yourself from non-linearities?
You can’t - but you can look for diagnostic defects
  Edges to structures
  Asymmetries in intensity
Test: reduce laser; does image reduce proportionately?
Avoid over-labeling
Avoid over-stimulating
“When in doubt, reduce intensity of stimulation”

## Slide 30: Microscope has non-linearities
light source
(image of arc)
Relay optics
Objective
Eyepieces
Camera

## Slide 31: Microscope has non-linearities
light source
(image of arc)
Relay optics
Objective
Eyepieces
Camera

## Slide 32: Objective lens better at collecting light near center

## Slide 33: Microscope has non-linearities
light source
(image of arc)
Relay optics
Objective
Eyepieces
Camera

## Slide 34: No free lunch from Confocal
Good: Single detector
Bad: Easy to saturate dye (less excitation, ISC)
Good: Single light source
Bad: Very sensitive to optical aberrations
Notes: Why is single detector good?

## Slide 35: Optical aberrations
Spherical aberration
Lateral chromatic aberration
Focus deeper below coverglass
Not corrected for
  spherical aberration
  chromatic aberration
Detector
Detector

## Slide 36: Solution: Flat Fielding (pixel by pixel correction)
Requirements:
Specimen of uniform intensity
Set of specimens of different known brightness
Slide with double-stick tape
Cut holes in tape
Drop dye in holes
Different [dye] in each hole
Coverslip over the top

## Slide 37
Intensity
[dye]
On pixel by pixel basis
Plot I vs dye concentration [dye]
Calculate slope & intercept
References:
Kindler & Kennedy (1996) J Neurosci Methods 68:61-70
Stollberg & Fraser (1988) J Cell Biol 107: 1397–1408.
Watch for warning sign
Sublinear -> self-quenching

## Slide 38: Flat-field correction of digital image
Corrected Image = [(Raw Image – Dark Frame) * M] / (Flat field Frame- Dark Frame)
Notes: http://www.olympusmicro.com/primer/digitalimaging/imageprocessingintro.html
http://www.princetoninstruments.com/cms/index.php/library/51-ccd-primer/152-flat-field-correction
http://www.roper.co.jp/Html/roper/tech_note/html/timagecal.htm

## Slide 39: So how many fluorophores does a given intensity equal?
J Neurosci Methods 105:55-63 (2001)

## Slide 40: Single molecule calibration: Beads with Ni-NTA; GFP::His6
0 pM
1 pM
10 pM
J Neurosci Methods 105:55-63 (2001)
Alternative: use viruses with defined numbers of GFP’s
Notes: Fig. 1. Images of single histidine-tagged green fluorescent protein (beads at increasing densities in an area of 150 × 150 pixels (16.75 × 16.75 µm). Panel A shows that the blank bead does not have the typical ‘GFP points’. Panels B– F show images of the beads after incubation with solutions of increasing concentrations (10− 12, 2.5 × 10− 12, 5 × 10− 12, 10− 11, and 2.5 × 10− 11 In panel B, the arrowheads point to the 38 images accepted as arising from GFP molecules.

## Slide 41: Macroscopic measurements based on single molecule calibration
Intensity proportional to # of GFP on beads over 3 orders of magnitude (Olympus)
Over 4 orders of magnitude for Nikon microscope
Density of membrane proteins can be measured to accuracy of about 20%
J Neurosci Methods 105:55-63 (2001)
Notes: Fig. 2. Calibrations of beads. These results were obtained using the Olympus microscope at an incident beam intensity of 22.8 W/cm2. (A) Histogram of fluorescence intensity distributions, fitted to a Gaussian with an average value of 376.4±2.6 counts/s (n=178). (B) The X axes are the [His6-GFP] in the incubation solutions, and the resulting expected density of green fluorescent protein (GFP) molecules on the beads. The expected density assumes that binding to the beads is quantitative and that all bound His6-GFP molecules are active, with the average fluorescence as in panel A. The Y value is the average fluorescence intensity, in counts/μm2. Results for beads carrying the highest six densities were multiplied by a ‘photobleaching corrective factor’<1 to correct for reduced bleaching due to the use of neutral density filters on the excitation beam (see text). The correction factor for the 2.0 OD filter (*) is 0.781, for the 1.0 OD filter (†) 0.803, and for the 0.7 OD filter (‡) 0.825. Standard errors of the observations are smaller than the symbol size. (C) Measured intensity is compared with the expected intensity, computed as expected density×average fluorescence counts per GFP.

## Slide 42: Multispectral Imaging
Instead of Z – stacks, collect λ – stacks
Spectral or Lambda Scanning

## Slide 43: Spectral image dataset
-stack
 can be:
(i) excitation
images acquired in a single channel at different excitation
(ii) emission
images acquired at a single excitation in several channels at different (emission)
Garini et al, Cytometry Part A, 2006

## Slide 44: Spectral image dataset
Garini et al, Cytometry Part A, 2006
Notes: Wave-length scans (A-D); Spatial Scan (E) is what will concentrate on.

## Slide 45: Spectral imaging methods: Spatial-scan
3 Different ways used by microscope companies
Notes: http://zeiss-campus.magnet.fsu.edu/articles/spectralimaging/introduction.html

## Slide 46: Dispersion through refraction versus diffraction
Diffraction grating
Refraction through prism
Note how longer wavelengths (red) diffract at greater angle than shorter wavelengths (blue) but they refract at smaller angle than shorter wavelengths.
Notes: Comparison of the spectra obtained from a diffraction grating by diffraction (1), and a prism by refraction (2). Longer wavelengths (red) are diffracted more, but refracted less than shorter wavelengths (violet).
"Comparison refraction diffraction spectra" by Cmglee - Own work. Licensed under CC BY-SA 3.0 via Wikimedia Commons - http://commons.wikimedia.org/wiki/File:Comparison_refraction_diffraction_spectra.svg#mediaviewer/File:Comparison_refraction_diffraction_spectra.svg

## Slide 47: Monochromator: Optical instrument for generating single colors
Used in optical measuring instruments
How a monochromator works according to the principle of dispersion
Most actually disperse through diffraction, not prism
Entrance Slit
Monochromator (Prism Type)
Exit Slit
Notes: A monochromator is an optical device that transmits a mechanically selectable narrow band of wavelengths of light or other radiation chosen from a wider range of wavelengths available at the input.

## Slide 48: Spectral imaging with a grating

## Slide 49: History of the Zeiss META detector
Where did the idea of a multichannel detector come from?

## Slide 50: History of the Zeiss META detector
Where did the idea of a multichannel detector come from?
Collaboration between the Jet Propulsion Laboratory, Scott Fraser’s lab here at Caltech and Zeiss

## Slide 51: Airborne Visible/Infrared Imaging Spectrometer (AVIRIS)
Instrument for earth imaging and ecological research.
Instrument has 224 detectors.
Covers a range from 380 nm to 2500 nm.
Notes: http://www.dfrc.nasa.gov/Gallery/Photo/SR-71/HTML/EC94-42883-2.html

## Slide 52: Airborne Visible/Infrared Imaging Spectrometer (AVIRIS)
Original
Next Generation (AVIRISng)
Notes: Note size of the detector. Too big for confocal.
http://airbornescience.jpl.nasa.gov/instruments/avirisng
http://aviris.jpl.nasa.gov/aviris/concept.html

## Slide 53: History of the Zeiss META detector
Zeiss META had 8* channel detector
Replaced by 32 channel Quasar detector
Notes: First Meta
http://bair.beatson.gla.ac.uk/pages1/confocals/SpectralUnmixing.htm
http://www.zeiss.com/microscopy/en_us/campus/spectral-imaging/practical-aspects-of-spectral-imaging.html#configuration

## Slide 54: Spectral imaging with a prism and mirrors

## Slide 55: Spectral image dataset
-stack
 can be:
(i) excitation
images acquired in a single channel at different excitation
(ii) emission
images acquired at a single excitation in several channels at different (emission)
Garini et al, Cytometry Part A, 2006
Notes: Leica feature

## Slide 56: Leica lambda squared map
White light laser that emits from 470 to 670 nanometers

## Slide 57: Choose spectrally well-separated dyes
Source: Zimmermann, T., 2005. Spectral Imaging and Linear Unmixing in Light Microscopy, in: Rietdorf, J. (Ed.), Microscopy Techniques. Springer Berlin Heidelberg, pp. 245-265.
if not possible: use spectral unmixing!
Notes: Wave-length scans (A-D);

## Slide 58: Spectral unmixing: general concept
Multi-channel Detector
Collect Lambda Stack
Raw Image
Unmixed Image
Derive Emission Fingerprints
FITC
Sytox-green
Notes: Important! For this experiment need 3 samples! One with each color alone then one with both. Crucial for distinguishing overlap.
http://bair.beatson.gla.ac.uk/pages1/confocals/SpectralUnmixing.htm

## Slide 59: Linear spectral unmixing: principle
To solve and obtain Ai for each pixel
From Michael Liebling, UCSB
Notes: Signal = relative contribution X reference spectra

## Slide 60: Linear spectral unmixing: principle
From Michael Liebling, UCSB
2 possibilities:

## Slide 61: Spectral unmixing
8 channel detector (can you guess the instrument used?)
Using Emission spectra
Example of parallel acquisition
Reference spectra important
Notes: Fig. 2 Spectral imaging of fluorescence signals.Contributions of CFP, GFP and YFP to eight
successive spectral channels are shown. The distribution of emission signal to the channels
is a direct representation of the fluorophore emission spectrum and constitutes a spectral
signature.With linear unmixing using these spectral signatures as reference, even combined
and mixed signals can be clearly separated into the fluorophores that contribute to the total
signal

## Slide 62: Spectral unmixing: GFP/YFP

## Slide 63: Spectral unmixing: Leica lambda stack

## Slide 64: Spectral unmixing: Leica lambda stack

## Slide 65: Spectral unmixing: Leica lambda stack

## Slide 66: Spectral unmixing of autofluorescence
Red and green arrows indicate regions from which sample spectra were obtained.
Blue = computed spectrum
(a) Image obtained at the peak of one of the quantum dots.
(b) Unmixed image of the 570-nm quantum dot.
(c) Unmixed image of the 620-nm quantum dot.
(d) Combined pseudocolor image of (b) (green), (c), and autofluorescence channel (in white).
Mansfield et al, Journal of Biomedical Optics (2005)

## Slide 67: Determine the two photon spectra of uncharacterized dye
In vivo Hair Cell Dye, FM1-43 Spectra

## Slide 68: Spectral or Lambda Scanning
Separate very similar colored fluorophores
  e.g. FITC and Sytox green.
Could be used to eliminate non-specific background fluorescence that has different emission spectra.
Different technologies for spectrum detection
  Sequentially (Leica SP)
  Simultaneously (Zeiss QUASAR)