## Slide 1: Biology 227: Methods in Modern Microscopy
Andres Collazo, Director Biological Imaging Facility
Chai Tong Ng, Graduate Student, TA
Week 10: Super-Resolution Microscopy

## Slide 2: Spatial Resolution of Biological Imaging Techniques
Resolution is diffraction limited.
Abbe (1873) reported that smallest resolvable distance between two points (d) using a conventional microscope may never be smaller than half the wavelength of the imaging light (~200 nm)
Ernst Abbe (1840-1905)
Notes: http://www.nature.com/milestones/milelight/full/milelight03.html

## Slide 3: Super-resolution microscopy
Many ways to achieve
Some more super than others.
Notes: They fall into two broad categories, "true" super-resolution techniques, which capture information contained in evanescent waves, and "functional" super-resolution techniques, which use clever experimental techniques and known limitations on the matter being imaged to reconstruct a super-resolution image.[2]
True subwavelength imaging techniques include those that utilize the Pendry Superlens and near field scanning optical microscopy, the 4Pi Microscope and structured illumination microscopy technologies like SIM and SMI. However, the majority of techniques of importance in biological imaging fall into the functional category.
There are two major groups of methods for functional super-resolution microscopy:
Deterministic super-resolution: The most commonly used emitters in biological microscopy, fluorophores, show a nonlinear response to excitation, and this nonlinear response can be exploited to enhance resolution. These methods include STED, GSD, RESOLFT and SSIM.
Stochastic super-resolution: The chemical complexity of many molecular light sources gives them a complex temporal behaviour, which can be used to make several close-by fluorophores emit light at separate times and thereby become resolvable in time. These methods include SOFI and all single-molecule localization methods (SMLM) such as SPDM, SPDMphymod, PALM, FPALM, STORM and dSTORM.

## Slide 4: Super-resolution microscopy
Could always do super-resolution if could label points with different colors
Separate with different fluorescent filters (spectral unmixing)
Why fluorescence is such an important illumination technique
Hell, S.W., 2009. Microscopy and its focal switch. Nat Meth 6, 24-32.
Notes: http://www.nature.com/nmeth/journal/v6/n1/full/nmeth.1291.html

## Slide 5: Spatial Resolution of Biological Imaging Techniques
Notes: http://zeiss-campus.magnet.fsu.edu/articles/superresolution/introduction.html

## Slide 6: Super-resolution microscopy
“True” super-resolution techniques
  Subwavelength imaging
  Capture information in evanescent waves
“Functional” super-resolution techniques
  Deterministic
    Exploit nonlinear responses of fluorophores
  Stochastic
    Exploit the complex temporal behaviors of fluorophores
Notes: They fall into two broad categories, "true" super-resolution techniques, which capture information contained in evanescent waves, and "functional" super-resolution techniques, which use clever experimental techniques and known limitations on the matter being imaged to reconstruct a super-resolution image.[2]
True subwavelength imaging techniques include those that utilize the Pendry Superlens and near field scanning optical microscopy, the 4Pi Microscope and structured illumination microscopy technologies like SIM and SMI. However, the majority of techniques of importance in biological imaging fall into the functional category.
There are two major groups of methods for functional super-resolution microscopy:
Deterministic super-resolution: The most commonly used emitters in biological microscopy, fluorophores, show a nonlinear response to excitation, and this nonlinear response can be exploited to enhance resolution. These methods include STED, GSD, RESOLFT and SSIM.
Stochastic super-resolution: The chemical complexity of many molecular light sources gives them a complex temporal behaviour, which can be used to make several close-by fluorophores emit light at separate times and thereby become resolvable in time. These methods include SOFI and all single-molecule localization methods (SMLM) such as SPDM, SPDMphymod, PALM, FPALM, STORM and dSTORM.

## Slide 7: Spatial Resolution of Biological Imaging Techniques
“True” super-resolution
“Functional”
Notes: http://zeiss-campus.magnet.fsu.edu/articles/superresolution/introduction.html

## Slide 8: “Functional” super-resolution techniques
Deterministic
  Reversible Saturable (or Switchable) OpticaL Fluorescence Transitions (RESOLFT)
    STimulated Emission Depletion (STED)
    Ground State Depletion (GSD)
Stochastic
  STochastic Optical Reconstruction Microscopy (STORM)
  Photo Activated Localization Microscopy (PALM)
  Fluorescence Photo-Activation Localization Microscopy (FPALM)
Notes: A trio of techniques independently surfaced during a four-month period in 2006 that elegantly demonstrated the sequential and stochastic readout with precise localization of multiple single molecules from specimens that were fluorescently labeled with photoswitchable dyes or optical highlighter fluorescent proteins at densities approaching those commonly employed in widefield and confocal fluorescence microscopy (ranging up to 100,000 molecules per square micrometer). These nearly identical single-molecule superresolution imaging techniques have been termed stochastic optical reconstruction microscopy (STORM; see Figure 1), photoactivated localization microscopy (PALM), and fluorescence photoactivation localization microscopy (FPALM). They rely on an activation laser applied at low power to trigger emission switching so that any particular molecule has a small probability of being photoactivated, but the majority of the population remains in the original (dark or native) emissive state. Although the major difference between the methods as they were originally published was simply the nature of the fluorescent probes used to achieve superresolution imaging, these techniques are virtually identical. In fact, the original STORM research report mentioned the use of fluorescent proteins, whereas the first PALM paper discussed the application of caged synthetic fluorophores.
http://www.microscopyu.com/articles/superresolution/stormintro.html
They fall into two broad categories, "true" super-resolution techniques, which capture information contained in evanescent waves, and "functional" super-resolution techniques, which use clever experimental techniques and known limitations on the matter being imaged to reconstruct a super-resolution image.[2]
True subwavelength imaging techniques include those that utilize the Pendry Superlens and near field scanning optical microscopy, the 4Pi Microscope and structured illumination microscopy technologies like SIM and SMI. However, the majority of techniques of importance in biological imaging fall into the functional category.
There are two major groups of methods for functional super-resolution microscopy:
Deterministic super-resolution: The most commonly used emitters in biological microscopy, fluorophores, show a nonlinear response to excitation, and this nonlinear response can be exploited to enhance resolution. These methods include STED, GSD, RESOLFT and SSIM.
Stochastic super-resolution: The chemical complexity of many molecular light sources gives them a complex temporal behaviour, which can be used to make several close-by fluorophores emit light at separate times and thereby become resolvable in time. These methods include SOFI and all single-molecule localization methods (SMLM) such as SPDM, SPDMphymod, PALM, FPALM, STORM and dSTORM.

## Slide 9: Reversible Saturable (or Switchable) Optical Fluorescence Transitions (RESOLFT)
Includes
STED
GSD
Notes: Extremely high intensities are necessary for switching off the excited singlet state with STED (using what is termed a depletion laser), but switching to a metastable triplet or similar dark state (GSD) requires a light intensity between a thousand and a million times lower.
http://zeiss-campus.magnet.fsu.edu/articles/superresolution/introduction.html
http://www.microscopyu.com/articles/superresolution/stormintro.html

## Slide 10: STED: STimulated Emission Depletion
http://zeiss-campus.magnet.fsu.edu

## Slide 11: STED Microscopy means scanning a smaller focal spot across the sample
Point Spread Function: Confocal vs STED
measured with fluorescent nanoparticles under the same conditions
y
x
y
x
Typical lateral (X-Y) resolution in a Confocal: 200x200 nm
Typical lateral (X-Y) FWHM (Full Width Half Maximum)
in STED is 90x90 nm
STED z resolution is confocal (500 nm)
STED enables separation of structures even smaller than its FWHM due to the sharp peak. Actual resolution is in the order of 70 nm for raw data (without deconvolution)

## Slide 12: “Functional” super-resolution techniques
Deterministic
  Reversible Saturable (or Switchable) OpticaL Fluorescence Transitions (RESOLFT)
    STimulated Emission Depletion (STED)
    Ground State Depletion (GSD)
Stochastic
  STochastic Optical Reconstruction Microscopy (STORM)
  Photo Activated Localization Microscopy (PALM)
  Fluorescence Photo-Activation Localization Microscopy (FPALM)
Notes: A trio of techniques independently surfaced during a four-month period in 2006 that elegantly demonstrated the sequential and stochastic readout with precise localization of multiple single molecules from specimens that were fluorescently labeled with photoswitchable dyes or optical highlighter fluorescent proteins at densities approaching those commonly employed in widefield and confocal fluorescence microscopy (ranging up to 100,000 molecules per square micrometer). These nearly identical single-molecule superresolution imaging techniques have been termed stochastic optical reconstruction microscopy (STORM; see Figure 1), photoactivated localization microscopy (PALM), and fluorescence photoactivation localization microscopy (FPALM). They rely on an activation laser applied at low power to trigger emission switching so that any particular molecule has a small probability of being photoactivated, but the majority of the population remains in the original (dark or native) emissive state. Although the major difference between the methods as they were originally published was simply the nature of the fluorescent probes used to achieve superresolution imaging, these techniques are virtually identical. In fact, the original STORM research report mentioned the use of fluorescent proteins, whereas the first PALM paper discussed the application of caged synthetic fluorophores.
http://www.microscopyu.com/articles/superresolution/stormintro.html
These three methods were published independently during a short period of time and their principle is identical. STORM was originally described using Cy5 and Cy3 dyes attached to nucleic acids or proteins,[47] while PALM and FPALM was described using photoswitchable fluorescent proteins.[48][49]
http://en.wikipedia.org/wiki/Super-resolution_microscopy#Stochastic_functional_techniques
They fall into two broad categories, "true" super-resolution techniques, which capture information contained in evanescent waves, and "functional" super-resolution techniques, which use clever experimental techniques and known limitations on the matter being imaged to reconstruct a super-resolution image.[2]
True subwavelength imaging techniques include those that utilize the Pendry Superlens and near field scanning optical microscopy, the 4Pi Microscope and structured illumination microscopy technologies like SIM and SMI. However, the majority of techniques of importance in biological imaging fall into the functional category.
There are two major groups of methods for functional super-resolution microscopy:
Deterministic super-resolution: The most commonly used emitters in biological microscopy, fluorophores, show a nonlinear response to excitation, and this nonlinear response can be exploited to enhance resolution. These methods include STED, GSD, RESOLFT and SSIM.
Stochastic super-resolution: The chemical complexity of many molecular light sources gives them a complex temporal behaviour, which can be used to make several close-by fluorophores emit light at separate times and thereby become resolvable in time. These methods include SOFI and all single-molecule localization methods (SMLM) such as SPDM, SPDMphymod, PALM, FPALM, STORM and dSTORM.

## Slide 13: Single-molecule localization (SML) microscopy Stochastic functional techniques
Notes: http://zeiss-campus.magnet.fsu.edu/articles/superresolution/palm/practicalaspects.html
http://www.microscopyu.com/articles/superresolution/stormintro.html

## Slide 14: Single-molecule localization microscopy Stochastic functional techniques
STED vs STORM
How STORM works

## Slide 15: Single-molecule localization microscopy
Must have sufficient density of molecules being localized
Notes: Among the critical aspects that must be considered are the accuracy of each individual localization measurement, the density of probes that have been localized in the final image (commonly termed molecular density), and the physical size of the labels themselves. The relationship between resolution and the localization precision of a single molecule is readily determined. The ability to resolve two fluorescent molecules as separate entities is limited by the localization precision, which determines the position (and uncertainty) of each molecule and thus the distance between the pair. The localization precision, in turn, is primarily dependent upon the number of photons collected from the fluorescent molecules during a single activation-deactivation cycle, provided the background noise is negligible.
Presented in Figure 7 is an important concept involving one of the most critical factors influencing spatial resolution in single-molecule superresolution imaging. The resolution versus molecular density is represented in Figure 7 as a series of yellow pixels in a test pattern. Features that might be present in a specimen are imaged at a progressively lower signal-to-noise ratio (fraction of pixels measured) as the molecular density decreases (from the bottom to the top of Figure 7) and these features become unresolvable when the mean molecular separation approaches feature size (the top row in Figure 7). Thus, as described in Equation (5), in order to achieve N-fold higher resolution in dimensions D, ND-fold more pixels must be acquired. In order to realize such a resolution gain without compromising either the imaging speed or the signal-to-noise ratio, the signal collection rate (photons detected per second) must increase by a factor of at least ND, which requires an ND-fold higher exposure to the excitation laser for each image required.

## Slide 16: Each super-resolution techniques have pluses and minuses but all methods are improving
Schermelleh, L., Heintzmann, R., Leonhardt, H., 2010. A guide to super-resolution fluorescence microscopy. The Journal of Cell Biology 190, 165-175.
Notes: http://jcb.rupress.org/content/190/2/165.full

## Slide 17: Super-resolution requirements
High power lasers
Special fluorophores
Concentration of fluorophores
Special optics
Computational processing
Fast detectors
Sensitive detectors
Precise X,Y,Z positioning

## Slide 18: Super-resolution microscopy
Functional super-resolution techniques won Nobel
SIM did not win
SIM not considered to be super-resolution
Notes: They fall into two broad categories, "true" super-resolution techniques, which capture information contained in evanescent waves, and "functional" super-resolution techniques, which use clever experimental techniques and known limitations on the matter being imaged to reconstruct a super-resolution image.[2]
True subwavelength imaging techniques include those that utilize the Pendry Superlens and near field scanning optical microscopy, the 4Pi Microscope and structured illumination microscopy technologies like SIM and SMI. However, the majority of techniques of importance in biological imaging fall into the functional category.
There are two major groups of methods for functional super-resolution microscopy:
Deterministic super-resolution: The most commonly used emitters in biological microscopy, fluorophores, show a nonlinear response to excitation, and this nonlinear response can be exploited to enhance resolution. These methods include STED, GSD, RESOLFT and SSIM.
Stochastic super-resolution: The chemical complexity of many molecular light sources gives them a complex temporal behaviour, which can be used to make several close-by fluorophores emit light at separate times and thereby become resolvable in time. These methods include SOFI and all single-molecule localization methods (SMLM) such as SPDM, SPDMphymod, PALM, FPALM, STORM and dSTORM.

## Slide 19: Structured Illumination Microscopy (SIM) Option for wide-field microscope
Zeiss Apotome 2
Keyence BZ-X700
Notes: Grab bag of methods that are not

## Slide 20: Structured Illumination - The Contrast is the Key
Position 1
Position 2
Position 3
Optical section
No visible grid  out of focus light
Visible grid  structures in focus
Notes: The algorithm is looking for a change in the local contrast. The algorithm is demodulating the images.
Simple image analysis algorithm detects the brightness differences of the 3 raw images and removes the out-of-focus signals in the image (demodulation)

## Slide 21: How does super-resolution SIM differ from normal SIM?
It’s all due to the number of different gratings
Zeiss Apotome uses 1 grating
Keyence uses up to 2 gratings
For super-resolution Mats Gustafsson found that 3 different grating orientations were required (gratings finer also)
Notes: Heintzmann, R., Cremer, C.G., 1999. Laterally modulated excitation microscopy: improvement of resolution by using a diffraction grating, pp. 185-196.
Gustafsson, M.G.L., 2000. Surpassing the lateral resolution limit by a factor of two using structured illumination microscopy. Journal of Microscopy 198, 82-87.
Fig. 1. Concept of resolution enhancement by structured illumination. (a) If two line patterns are superposed (multiplied), their product will
contain moireÂ fringes (seen here as the apparent vertical stripes in the overlap region). (b) A conventional microscope is limited by diffraction.
The set of low-resolution information that it can detect defines a circular `observable region' of reciprocal space. (c) A sinusoidally striped
illumination pattern has only three Fourier components. The possible positions of the two side components are limited by the same circle that
defines the observable region (dashed). If the sample is illuminated with such structured light, moireÂ fringes will appear which represent
information that has changed position in reciprocal space. The amounts of that movement correspond to the three Fourier components of the
illumination. The observable region will thus contain, in addition to the normal information, moved information that originates in two offset
regions (d). From a sequence of such images with different orientation and phase of the pattern, it is possible to recover information from an
area twice the size of the normally observable region, corresponding to twice the normal resolution (e).

## Slide 22: Examples of SR-SIM images from Gustafsson
121 nm beads
Actin cytoskeleton
Notes: Fig. 3. A cluster of fluorescent microspheres of nominal diameter
121 nm, as imaged by conventional (a), confocal (b), and
structured illumination (c) microscopy. The confocal microscope
was operated under maximum resolution conditions. The apparent
size (FWHM) of isolated beads is approximately 130 nm in the
structured illumination image, compared to 290 and 210 nm for
conventional and confocal microscopy.
Fig. 4. The actin cytoskeleton at the edge of a HeLa cell, as imaged by (a, c) conventional and (b, d) structured illumination microscopy. (c, d)
Enlargements of the boxed areas in (a) and (b), respectively. Fibres separated by less than the resolution limit of the conventional microscope
are well resolved using structured illumination (d). The apparent widths (FWHM) of the finest protruding fibres [small arrows in (a, b)] are
lowered to 110±120 nm in (b), compared to 280±300 nm in (a).

## Slide 23: 3D SR-SIM uses SLM from Gustafsson (2011)
MitoTracker Green acquired at 20 sec intervals
Notes: Figure 2 | Live 3D SIM and conventional wide-field microscopy images of a HeLa cell stained with MitoTracker Green. (a) Maximum-intensity projection along z dimension through the cell volume at time 0 of a 50-time-point series. Each SIM volume was acquired in 20.33 s (that is, 35-ms exposure time × 38 axial planes × 15 patterns + 38 axial planes × 10-ms z-stage settle time). (b) One x–z cross-section of the same volume sliced through the dashed line shown in a. (c–f) Single-plane x–y slices corresponding to the boxed regions in a. (g) Eight time frames of the region boxed with dashed line in a. Each frame is a maximum-intensity projection along z over a 1.3 μm thickness that contains the featured ‘Y’-shaped mitochondrion. Scale bars, 2 μm (a,b) and 1 μm (c–g).
Shao, L., Kner, P., Rego, E.H., Gustafsson, M.G.L., 2011. Super-resolution 3D microscopy of live whole cells using structured illumination. Nat Meth 8, 1044-1046.

## Slide 24: Schermelleh, L., Heintzmann, R., Leonhardt, H., 2010. A guide to super-resolution fluorescence microscopy. The Journal of Cell Biology 190, 165-175.
Notes: http://jcb.rupress.org/content/190/2/165.full

## Slide 25: Zeiss Airyscan yet another means to achieve super-resolution
Available here at Caltech in the BIF

## Slide 26: Conventional scanning confocal and the 1 A.u. limit. PH = 1 A.u. is the standard setting for resolution.
A point-like emitter generates a
diffraction limited pattern (~ PSF)
PH = 1 A.u.
scan
scan
Intensity
excitation
At 1 A.u. The PSF is mapped directly 1:1
excitation and detection
are scanned in sync
~ 240 nm
detection
Notes: The practical limit for the pinhole-size in confocal microscopy is 1.A.u. the resolution achieved at that setting is comparable to the classical limits postulated by Abbé or Rayleigh. The spatial image of a point-like emitter recorded with a confocal microscope with a 1 A.u. pinhole will correspond to the classically defined PSF.

## Slide 27: Scanning confocal and the 1 Airy unit pinhole “limit”. At PH < 1 AU signal loss is larger than gain in resolution.
The potential to increase resolution by closing the pinhole (PH) is not a new idea (red)
One finds this idea described in many places, including Pawley’s Handbook of Confocal Microscopy
The price to pay, however, is
sacrificing 95% of the signal
(green).
The famous 1 AU setting
is more of a practical barrier
than a theoretical limit.
Fortunately, now there is
a clever workaround with
Airyscan...
Notes: One main aspect of Airyscan is its capability to achieve superresolution (in the sense of Abbé or Rayleigh). The idea, indeed, is not new – It is widely known that confocal microscopes carry the capability to convolve excitation with detection and result in a factor-sqrt(2) ammeliorated resolution. However, in order to tap into this regime, the pinhole must be much smaller than the PSF itself. The loss of signal in this case poses, in the mojority of cases, an unsurmountable difficulty. As noted previously, Airyscan acts as if scanning with 0.2 A.u. pinholes, however, because the light is collected from the whole array, each pixel is in effect scanned at least 32 times. Since all photons are collected the loss of signal is virtually avoided.

## Slide 28: Airyscan: PH ~ 0.2 A.u. scanning without loss. A single element improves resolution
detection
PH = 0.2 a.u.
Intensity
excitation
A point-like emitter generates a
diffraction limited pattern (~ PSF)
scan
scan
By collecting just the central element
the PSF is weaker but smaller
~170nm
Notes: Airyscan acts as if scanning with 0.2 A.u. pinholes, however, because the light is collected from the whole array, each pixel is - in effect - scanned at least 32 times. Since all photons are collected the loss of signal is virtually avoided. The PSF can be re-mapped as shown, generating a smaller PSF by a factor sqrt(2). This is true in every direction (including z – not shown, here). At this level of oversampling, one is generating ideal conditions for deconvolution. A deconvolution step applied to these images generates a further increase, of about 1.2-fold, in resolution. This additional push in resolution, however, depends on image quality and cannot be expected to work perfectly all the times.

## Slide 29: Airyscan Detection Unique 32 GaAsP-PMT design
32 GaAsP-PMT array
adaptive zoom optics
filter wheel
LSM 880 output port
X
Y
Replaced with New Design
Notes: Airyscan is a 32-channel GaAsP detector array arranged in a hexagonal grid. The full intensity distribution passing the pinhole reaches this detector. Each GaAsP subunit can be seen as a detector with a smaller pinhole. In effect every pixel on the image is generated from 32 subunits – either using the space information to generate superresolution or collecting all photons as an ordinary detector...

## Slide 30: Airyscan: PH ~ 0.2 A.u. scanning without loss. A single element improves resolution
detection
PH = 1,25 a.u.
Intensity
excitation
A point-like emitter generates a
diffraction limited pattern (~ PSF)
scan
scan
subunit
~ 0.2 A.u.
By collecting just the central element
the PSF is weaker but smaller
~170nm
Notes: Airyscan acts as if scanning with 0.2 A.u. pinholes, however, because the light is collected from the whole array, each pixel is - in effect - scanned at least 32 times. Since all photons are collected the loss of signal is virtually avoided. The PSF can be re-mapped as shown, generating a smaller PSF by a factor sqrt(2). This is true in every direction (including z – not shown, here). At this level of oversampling, one is generating ideal conditions for deconvolution. A deconvolution step applied to these images generates a further increase, of about 1.2-fold, in resolution. This additional push in resolution, however, depends on image quality and cannot be expected to work perfectly all the times.

## Slide 31: Airyscan: PH ~ 0.2 A.u. scanning without loss. A single element – even if offset still improves resolution
subunit
~ 0.2 A.u.
detection
PH = 1,25 a.u.
Intensity
excitation
A point-like emitter generates a
diffraction limited pattern (~ PSF)
scan
scan
By using an element offset from the
center, resolution is still improved
Notes: Airyscan acts as if scanning with 0.2 A.u. pinholes, however, because the light is collected from the whole array, each pixel is - in effect - scanned at least 32 times. Since all photons are collected the loss of signal is virtually avoided. The PSF can be re-mapped as shown, generating a smaller PSF by a factor sqrt(2). This is true in every direction (including z – not shown, here). At this level of oversampling, one is generating ideal conditions for deconvolution. A deconvolution step applied to these images generates a further increase, of about 1.2-fold, in resolution. This additional push in resolution, however, depends on image quality and cannot be expected to work perfectly all the times.

## Slide 32: Airyscan: PH ~ 0.2 A.u. scanning without loss. Combining the data
subunit
~ 0.2 A.u.
detection
PH = 1,25 a.u.
Intensity
excitation
scan
scan
Image formation:
1 = Offset signal
2 = Add signal
Notes: Airyscan acts as if scanning with 0.2 A.u. pinholes, however, because the light is collected from the whole array, each pixel is - in effect - scanned at least 32 times. Since all photons are collected the loss of signal is virtually avoided. The PSF can be re-mapped as shown, generating a smaller PSF by a factor sqrt(2). This is true in every direction (including z – not shown, here). At this level of oversampling, one is generating ideal conditions for deconvolution. A deconvolution step applied to these images generates a further increase, of about 1.2-fold, in resolution. This additional push in resolution, however, depends on image quality and cannot be expected to work perfectly all the times.

## Slide 33: Airyscan: PH ~ 0.2 A.u. scanning without loss. With 32 detectors all photons are recorded and mapped at the same time
detection
PH = 1,25 a.u.
Intensity
excitation
scan
Get an extra push beyond the limit using deconvolution (down to 1:1,7)
~ 140 nm
subunit
~ 0.2 A.u.
All elements are acquired simultaneously, and can be remapped for better resolution and sensitivity
The PSF is mapped directly 1:√2
Notes: Airyscan acts as if scanning with 0.2 A.u. pinholes, however, because the light is collected from the whole array, each pixel is - in effect - scanned at least 32 times. Since all photons are collected the loss of signal is virtually avoided. The PSF can be re-mapped as shown, generating a smaller PSF by a factor sqrt(2). This is true in every direction (including z – not shown, here). At this level of oversampling, one is generating ideal conditions for deconvolution. A deconvolution step applied to these images generates a further increase, of about 1.2-fold, in resolution. This additional push in resolution, however, depends on image quality and cannot be expected to work perfectly all the times.

## Slide 34: TQFR survey
Teaching Quality Feedback Report (TQFR) survey period for WI 2017-18 will open today Monday, March 5, 2018
Link is here: https://access.caltech.edu/tqfr/taker/queue